## Energy Calibration
isotope | peak | literature energy /keV | uncertainty | photons per 100 disintegrations | uncertainty | mean /keV | standard deviation | amplitude | mx + c | residual
Na22 | 0 | 511 | 0 | 180.7 | 0.2 | 511.865 | 1.715 | 28.194 | 511.8806 | -0.015631 |  | m | 0.99916 | -0.45152 | c | m | 1.0008 | 0.45215 | c
Na22 | 1 | 1274.537 | 0.007 | 99.94 | 0.13 | 1276.016 | 1.521 | 7.823 | 1276.0579 | -0.041874 |  | unc | 0.00011919 | 0.11573 | unc | unc | 0.00011939 | 0.11578 | unc
Na22 | 2 | 1785.537 | 0.007 |  |  | 1787.192 | 2.09 | 0.109 | 1787.4864 | -0.29436 |  | r^2 | 1 | 0.24202 |  | r^2 | 1 | 0.24223
Co60 | 0 | 1173.228 | 0.003 | 99.85 | 0.03 | 1174.989 | 1.43 | 0.385 | 1174.6639 | 0.32508 |  |  | 70272070.9511 | 20 |  |  | 70272070.9511 | 20
Co60 | 1 | 1332.492 | 0.004 | 99.9826 | 0.0006 | 1334.42 | 1.47 | 0.339 | 1334.0615 | 0.35853 |  |  | 4116164.882 | 1.1715 |  |  | 4123071.9407 | 1.1735
Cs137 | 0 | 661.657 | 0.003 | 84.99 | 0.2 | 662.744 | 1.265 | 50.483 | 662.664 | 0.080039
Eu152 | 4 | 121.7817 | 0.0003 | 28.41 | 0.13 | 122.204 | 1.131 | 23.384 | 122.336 | -0.13196 | Energy Calibration = |  | 0.99916(12) x - 0.45(12)
Eu152 | 6 | 244.6974 | 0.0008 | 7.55 | 0.04 | 245.386 | 1.163 | 3.168 | 245.3547 | 0.031269
Eu152 | 7 | 344.2785 | 0.0012 | 26.59 | 0.12 | 345.093 | 1.186 | 8.741 | 345.0193 | 0.073668
Eu152 | 8 | 367.7891 | 0.002 | 0.862 | 0.005 | 368.273 | 1.432 | 0.291 | 368.5496 | -0.27665
Eu152 | 9 | 411.1165 | 0.0012 | 2.238 | 0.01 | 412.011 | 1.232 | 0.563 | 411.9134 | 0.097623
Eu152 | 10 | 443.965 | 0.003 | 2.8 | 0.02 | 444.875 | 1.23 | 0.684 | 444.7894 | 0.085579
Eu152 | 12 | 688.67 | 0.005 | 0.841 | 0.006 | 689.813 | 1.1 | 0.143 | 689.6996 | 0.11339
Eu152 | 13 | 778.9045 | 0.0024 | 12.97 | 0.06 | 780.074 | 1.306 | 1.672 | 780.0098 | 0.064225
Eu152 | 14 | 867.38 | 0.003 | 4.243 | 0.023 | 868.646 | 1.353 | 0.426 | 868.5595 | 0.086536
Eu152 | 15 | 964.079 | 0.018 | 14.5 | 0.06 | 965.35 | 1.424 | 1.357 | 965.3395 | 0.010452
Eu152 | 16 | 1005.272 | 0.017 | 0.655 | 0.023 | 1005.753 | 1.41 | 0.131 | 1006.5671 | -0.81409
Eu152 | 17 | 1085.837 | 0.01 | 10.13 | 0.06 | 1087.382 | 1.653 | 0.891 | 1087.1996 | 0.18236
Eu152 | 18 | 1112.076 | 0.003 | 13.41 | 0.06 | 1113.482 | 1.445 | 1.104 | 1113.4606 | 0.021353
Eu152 | 20 | 1212.948 | 0.011 | 1.416 | 0.009 | 1214.423 | 1.31 | 0.098 | 1214.4172 | 0.0057698
Eu152 | 21 | 1299.142 | 0.008 | 1.633 | 0.009 | 1300.652 | 1.54 | 0.112 | 1300.6835 | -0.031506
Eu152 | 22 | 1408.013 | 0.003 | 20.85 | 0.08 | 1409.716 | 1.594 | 1.236 | 1409.6458 | 0.070203

## Energy Calibration (2)
isotope | peak | literature energy /keV | uncertainty | photons per 100 disintegrations | uncertainty | mean /keV | standard deviation | amplitude | unc | mx + c | residual
Na22 | 0 | 511 | 0 | 180.7 | 0.2 | 511.865 | 1.715 | 28.194 | 0.15578 | 512.0139 | -0.14894 | m | 0.99919 | -0.59677 | c | m | 1.0008 | 0.59793 | c
Na22 | 1 | 1274.537 | 0.007 | 99.94 | 0.13 | 1276.016 | 1.521 | 7.823 | 0.2785 | 1276.1725 | -0.15653 | unc | 0.00017345 | 0.16084 | unc | unc | 0.00017373 | 0.16089 | unc
Na22 | 2 | 1785.537 | 0.007 |  |  | 1787.192 | 2.09 | 0.109 | 2.7658 | 1787.5885 | -0.39653 | r^2 | 1 | 0.44899 |  | r^2 | 1 | 0.44936
Co60 | 0 | 1173.228 | 0.003 | 99.85 | 0.03 | 1174.989 | 1.43 | 0.385 | 1.2173 | 1174.7811 | 0.20795 |  | 33184927.5995 | 28 |  |  | 33184927.5995 | 28
Co60 | 1 | 1332.492 | 0.004 | 99.9826 | 0.0006 | 1334.42 | 1.47 | 0.339 | 1.3153 | 1334.1747 | 0.24529 |  | 6689806.7879 | 5.6446 |  |  | 6700709.1407 | 5.6538
Cs137 | 0 | 661.657 | 0.003 | 84.99 | 0.2 | 662.744 | 1.265 | 50.483 | 0.099983 | 662.7936 | -0.049585
Eu152 | 4 | 121.7817 | 0.0003 | 28.41 | 0.13 | 122.204 | 1.131 | 23.384 | 0.13891 | 122.4788 | -0.27478
Eu152 | 6 | 244.6974 | 0.0008 | 7.55 | 0.04 | 245.386 | 1.163 | 3.168 | 0.38269 | 245.4945 | -0.10854
Eu152 | 7 | 344.2785 | 0.0012 | 26.59 | 0.12 | 345.093 | 1.186 | 8.741 | 0.23266 | 345.1567 | -0.063709
Eu152 | 8 | 367.7891 | 0.002 | 0.862 | 0.005 | 368.273 | 1.432 | 0.291 | 1.4011 | 368.6864 | -0.41345
Eu152 | 9 | 411.1165 | 0.0012 | 2.238 | 0.01 | 412.011 | 1.232 | 0.563 | 0.93434 | 412.0491 | -0.038121
Eu152 | 10 | 443.965 | 0.003 | 2.8 | 0.02 | 444.875 | 1.23 | 0.684 | 0.84699 | 444.9244 | -0.049363
Eu152 | 12 | 688.67 | 0.005 | 0.841 | 0.006 | 689.813 | 1.1 | 0.143 | 1.7518 | 689.8286 | -0.015576
Eu152 | 13 | 778.9045 | 0.0024 | 12.97 | 0.06 | 780.074 | 1.306 | 1.672 | 0.55822 | 780.1365 | -0.062535
Eu152 | 14 | 867.38 | 0.003 | 4.243 | 0.023 | 868.646 | 1.353 | 0.426 | 1.1256 | 868.6841 | -0.038062
Eu152 | 15 | 964.079 | 0.018 | 14.5 | 0.06 | 965.35 | 1.424 | 1.357 | 0.64702 | 965.4618 | -0.11178
Eu152 | 16 | 1005.272 | 0.017 | 0.655 | 0.023 | 1005.753 | 1.41 | 0.131 | 2.0722 | 1006.6883 | -0.93532
Eu152 | 17 | 1085.837 | 0.01 | 10.13 | 0.06 | 1087.382 | 1.653 | 0.891 | 0.86031 | 1087.3189 | 0.063093
Eu152 | 18 | 1112.076 | 0.003 | 13.41 | 0.06 | 1113.482 | 1.445 | 1.104 | 0.72261 | 1113.5793 | -0.097268
Eu152 | 20 | 1212.948 | 0.011 | 1.416 | 0.009 | 1214.423 | 1.31 | 0.098 | 2.3093 | 1214.5334 | -0.11039
Eu152 | 21 | 1299.142 | 0.008 | 1.633 | 0.009 | 1300.652 | 1.54 | 0.112 | 2.3421 | 1300.7976 | -0.14556
Eu152 | 22 | 1408.013 | 0.003 | 20.85 | 0.08 | 1409.716 | 1.594 | 1.236 | 0.71728 | 1409.7572 | -0.041189
Ba133 | 5 | 80.9979 | 0.0011 | 33.31 | 0.3 | 81.328 | 1.177 | 2.9949 | 0.39596 | 81.6618 | -0.33377
Ba133 | 28 | 276.3989 | 0.0012 | 7.13 | 0.06 | 277.521 | 1.174 | 0.23223 | 1.4201 | 277.2218 | 0.29915
Ba133 | 29 | 302.8508 | 0.0005 | 18.31 | 0.11 | 304.051 | 1.188 | 0.55438 | 0.92461 | 303.6953 | 0.35572
Ba133 | 32 | 356.0129 | 0.0007 | 62.05 | 0.19 | 357.34 | 1.201 | 1.6905 | 0.53238 | 356.9007 | 0.43934
Ba133 | 35 | 383.8485 | 0.0012 | 8.94 | 0.06 | 386.33 | 2.042 | 0.26699 | 1.7468 | 384.7589 | 1.5711
Bi207 | 7 | 569.698 | 0.002 | 97.76 | 0.03 | 570.223 | 1.702 | 8.266 | 0.28661 | 570.7597 | -0.53672
Bi207 | 9 | 1063.656 | 0.003 | 74.58 | 0.22 | 1065.139 | 1.925 | 3.2037 | 0.4896 | 1065.1198 | 0.01915
Bi207 | 11 | 1770.228 | 0.009 | 6.871 | 0.026 | 1773.194 | 2.29 | 0.15869 | 2.3994 | 1772.2671 | 0.92693

## Energy Efficiency Calibration
| isotope | peak |  | literature energy /keV | uncertainty | photons per 100 disintegrations | uncertainty |  | mean /keV | standard deviation | amplitude | unc (in all 3 (min dp calc'd to)) |  | current activity /kBq | uncertainty | expected area | uncertainty |  | measured area | unc |  | ratio | unc | y | 1/x |  |  |  |  | fit | residual | fit line x | fit line y
MORE | Ba133 | 1 |  |  |  | 96.2 | 0.80623 |  | 30.948 | 1.07 | 17.0723 | 0.0005 |  | 9.403 | 0.58338 | 9045.7337 | 566.3086 |  | 45.7895 | 0.021439 |  | 0.005062 | 0.00031692 | 0.005062 | 0.032312 |  | 0.2 |  |  | 0.19193 | -0.18686 | 36 | #NUM!
PEAKS | Ba133 | 5 |  | 80.9979 | 0.0011 | 33.31 | 0.3 |  | 81.247 | 1.209 | 6.7974 | 0.0005 |  | 9.403 | 0.58338 | 3132.1558 | 196.3607 |  | 20.5997 | 0.008653 |  | 0.0065768 | 0.00041232 | 0.0065768 | 0.012308 |  |  |  |  | 0.073107 | -0.06653 | 41 | #NUM!
ARE | Ba133 | 20 |  | 276.3989 | 0.0012 | 7.13 | 0.06 |  | 277.546 | 1.201 | 0.55128 | 0.0005 |  | 9.403 | 0.58338 | 670.4374 | 41.9759 |  | 1.6596 | 0.0016562 |  | 0.0024754 | 0.00015501 | 0.0024754 | 0.003603 | m | 5.9397 | 0 | c | 0.021401 | -0.018926 | 46 | #NUM!
IN | Ba133 | 21 |  | 302.8508 | 0.0005 | 18.31 | 0.11 |  | 304.1 | 1.214 | 1.3228 | 0.0005 |  | 9.403 | 0.58338 | 1721.6984 | 107.3165 |  | 4.0254 | 0.0022503 |  | 0.002338 | 0.00014574 | 0.002338 | 0.0032884 | unc | 0.25072 | #N/A | unc | 0.019532 | -0.017194 | 51 | #NUM!
OUTPUT | Ba133 | 23 |  | 356.0129 | 0.0007 | 62.05 | 0.19 |  | 357.411 | 1.223 | 3.9353 | 0.0005 |  | 9.403 | 0.58338 | 5834.5923 | 362.4278 |  | 12.0641 | 0.0051649 |  | 0.0020677 | 0.00012844 | 0.0020677 | 0.0027979 | r^2 | 0.96559 | 0.0028458 |  | 0.016619 | -0.014551 | 56 | #NUM!
FILE | Ba133 | 24 |  | 383.8485 | 0.0012 | 8.94 | 0.06 |  | 385.95 | 1.86 | 0.54706 | 0.0005 |  | 9.403 | 0.58338 | 840.6326 | 52.4584 |  | 2.5506 | 0.0024299 |  | 0.0030341 | 0.00018936 | 0.0030341 | 0.002591 |  | 561.2581 | 20 |  | 0.01539 | -0.012356 | 61 | #NUM!
| Eu152 | 0 |  |  |  | 58.5 | 0.5831 |  | 39.908 | 1.136 | 49.1249 | 0.0005 |  | 5.2714 | 0.15293 | 3083.753 | 94.5988 |  | 139.8847 | 0.061585 |  | 0.045362 | 0.0013917 | 0.045362 | 0.025058 |  | 0.0045453 | 0.00016197 |  | 0.14884 | -0.10347 | 66 | #NUM!
| Eu152 | 1 |  |  |  | 14.82 | 0.20616 |  | 45.538 | 1.127 | 10.4524 | 0.0005 |  | 5.2714 | 0.15293 | 781.2174 | 25.1354 |  | 29.5277 | 0.013176 |  | 0.037797 | 0.0012162 | 0.037797 | 0.02196 |  |  |  |  | 0.13043 | -0.092638 | 71 | #NUM!
| Co60 | 0 |  | 1173.228 | 0.003 | 99.85 | 0.03 |  | 1176.789 | 1.648 | 19.3525 | 0.0005 |  | 10.0749 | 0.72812 | 10059.8292 | 727.033 |  | 79.9438 | 0.024343 |  | 0.0079468 | 0.00057433 | 0.0079468 | 0.00085107 |  |  |  |  | 0.0050551 | 0.0028917 | 76 | #NUM!
| Co60 | 1 |  | 1332.492 | 0.004 | 99.9826 | 0.0006 |  | 1336.513 | 1.704 | 16.7377 | 0.0005 |  | 10.0749 | 0.72812 | 10073.1886 | 727.9922 |  | 71.4915 | 0.021086 |  | 0.0070972 | 0.00051292 | 0.0070972 | 0.00074939 |  |  |  |  | 0.0044512 | 0.002646 | 81 | #NUM!
| Cs137 | 0 |  | 661.657 | 0.003 | 84.99 | 0.2 |  | 662.744 | 1.265 | 50.483 | 0.0005 |  | 15.2714 | 0.93406 | 12979.1607 | 794.4413 |  | 160.0758 | 0.063291 |  | 0.012333 | 0.00075492 | 0.012333 | 0.0015089 |  |  |  |  | 0.0089623 | 0.0033709 | 86 | #NUM!
| Eu152 | 4 |  | 121.7817 | 0.0003 | 28.41 | 0.13 |  | 122.204 | 1.131 | 23.384 | 0.0005 |  | 5.2714 | 0.15293 | 1497.597 | 43.9855 |  | 66.2936 | 0.029342 |  | 0.044267 | 0.0013003 | 0.044267 | 0.008183 |  |  |  |  | 0.048605 | -0.0043385 | 91 | #NUM!
| Eu152 | 6 |  | 244.6974 | 0.0008 | 7.55 | 0.04 |  | 245.386 | 1.163 | 3.168 | 0.0005 |  | 5.2714 | 0.15293 | 397.9886 | 11.7374 |  | 9.2354 | 0.0042296 |  | 0.023205 | 0.00068444 | 0.023205 | 0.0040752 |  |  |  |  | 0.024206 | -0.0010006 | 96 | #NUM!
| Eu152 | 7 |  | 344.2785 | 0.0012 | 26.59 | 0.12 |  | 345.093 | 1.186 | 8.741 | 0.0005 |  | 5.2714 | 0.15293 | 1401.658 | 41.154 |  | 25.9858 | 0.011056 |  | 0.018539 | 0.00054439 | 0.018539 | 0.0028978 |  |  |  |  | 0.017212 | 0.0013273 | 101 | #NUM!
| Eu152 | 8 |  | 367.7891 | 0.002 | 0.862 | 0.005 |  | 368.273 | 1.432 | 0.291 | 0.0005 |  | 5.2714 | 0.15293 | 45.4392 | 1.3444 |  | 1.0445 | 0.0018314 |  | 0.022988 | 0.00068131 | 0.022988 | 0.0027154 |  |  |  |  | 0.016129 | 0.006859 | 106 | #NUM!
| Eu152 | 9 |  | 411.1165 | 0.0012 | 2.238 | 0.01 |  | 412.011 | 1.232 | 0.563 | 0.0005 |  | 5.2714 | 0.15293 | 117.9733 | 3.463 |  | 1.7386 | 0.0016977 |  | 0.014738 | 0.00043285 | 0.014738 | 0.0024271 |  |  |  |  | 0.014416 | 0.00032108 | 111 | #NUM!
| Eu152 | 10 |  | 443.965 | 0.003 | 2.8 | 0.02 |  | 444.875 | 1.23 | 0.684 | 0.0005 |  | 5.2714 | 0.15293 | 147.5984 | 4.41 |  | 2.1089 | 0.0017639 |  | 0.014288 | 0.00042707 | 0.014288 | 0.0022478 |  |  |  |  | 0.013351 | 0.00093645 | 116 | #NUM!
| Eu152 | 12 |  | 688.67 | 0.005 | 0.841 | 0.006 |  | 689.813 | 1.1 | 0.143 | 0.0005 |  | 5.2714 | 0.15293 | 44.3322 | 1.3245 |  | 0.39429 | 0.0013902 |  | 0.008894 | 0.00026757 | 0.008894 | 0.0014497 |  |  |  |  | 0.0086107 | 0.00028338 | 121 | #NUM!
| Eu152 | 13 |  | 778.9045 | 0.0024 | 12.97 | 0.06 |  | 780.074 | 1.306 | 1.672 | 0.0005 |  | 5.2714 | 0.15293 | 683.697 | 20.086 |  | 5.4736 | 0.002659 |  | 0.0080058 | 0.00023523 | 0.0080058 | 0.0012819 |  |  |  |  | 0.0076143 | 0.00039149 | 126 | #NUM!
| Eu152 | 14 |  | 867.38 | 0.003 | 4.243 | 0.023 |  | 868.646 | 1.353 | 0.426 | 0.0005 |  | 5.2714 | 0.15293 | 223.6643 | 6.6013 |  | 1.4448 | 0.0017778 |  | 0.0064595 | 0.00019081 | 0.0064595 | 0.0011512 |  |  |  |  | 0.0068379 | -0.00037841 | 131 | #NUM!
| Eu152 | 15 |  | 964.079 | 0.018 | 14.5 | 0.06 |  | 965.35 | 1.424 | 1.357 | 0.0005 |  | 5.2714 | 0.15293 | 764.349 | 22.3998 |  | 4.8437 | 0.0024653 |  | 0.0063371 | 0.00018574 | 0.0063371 | 0.0010359 |  |  |  |  | 0.0061529 | 0.00018412 | 136 | #NUM!
| Eu152 | 16 |  | 1005.272 | 0.017 | 0.655 | 0.023 |  | 1005.753 | 1.41 | 0.131 | 0.0005 |  | 5.2714 | 0.15293 | 34.5275 | 1.5727 |  | 0.463 | 0.0017748 |  | 0.01341 | 0.00061295 | 0.01341 | 0.00099428 |  |  |  |  | 0.0059058 | 0.0075038 | 141 | #NUM!
| Eu152 | 17 |  | 1085.837 | 0.01 | 10.13 | 0.06 |  | 1087.382 | 1.653 | 0.891 | 0.0005 |  | 5.2714 | 0.15293 | 533.99 | 15.8117 |  | 3.6918 | 0.0023535 |  | 0.0069136 | 0.00020476 | 0.0069136 | 0.00091964 |  |  |  |  | 0.0054624 | 0.0014512 | 146 | #NUM!
| Eu152 | 18 |  | 1112.076 | 0.003 | 13.41 | 0.06 |  | 1113.482 | 1.445 | 1.104 | 0.0005 |  | 5.2714 | 0.15293 | 706.8911 | 20.7508 |  | 3.9988 | 0.0022791 |  | 0.0056568 | 0.00016609 | 0.0056568 | 0.00089808 |  |  |  |  | 0.0053344 | 0.00032246 | 151 | #NUM!
| Eu152 | 20 |  | 1212.948 | 0.011 | 1.416 | 0.009 |  | 1214.423 | 1.31 | 0.098 | 0.0005 |  | 5.2714 | 0.15293 | 74.6426 | 2.2169 |  | 0.3218 | 0.0016464 |  | 0.0043112 | 0.00012993 | 0.0043112 | 0.00082344 |  |  |  |  | 0.004891 | -0.00057978 | 156 | #NUM!
| Eu152 | 21 |  | 1299.142 | 0.008 | 1.633 | 0.009 |  | 1300.652 | 1.54 | 0.112 | 0.0005 |  | 5.2714 | 0.15293 | 86.0815 | 2.5421 |  | 0.43234 | 0.0019352 |  | 0.0050225 | 0.00015001 | 0.0050225 | 0.00076885 |  |  |  |  | 0.0045667 | 0.00045574 | 161 | #NUM!
| Eu152 | 22 |  | 1408.013 | 0.003 | 20.85 | 0.08 |  | 1409.716 | 1.594 | 1.236 | 0.0005 |  | 5.2714 | 0.15293 | 1099.0812 | 32.1643 |  | 4.9385 | 0.002528 |  | 0.0044933 | 0.00013152 | 0.0044933 | 0.00070936 |  |  |  |  | 0.0042134 | 0.00027988 | 166 | #NUM!
| Na22 | 0 |  | 511 | 0 | 180.7 | 0.2 |  | 511.865 | 1.715 | 28.194 | 0.0005 |  | 7.2884 | 1.0144 | 13170.1062 | 1833.0463 |  | 121.2023 | 0.035401 |  | 0.0092028 | 0.0012809 | 0.0092028 | 0.0019536 |  |  |  |  | 0.011604 | -0.0024013 | 171 | #NUM!
| Na22 | 1 |  | 1274.537 | 0.007 | 99.94 | 0.13 |  | 1276.016 | 1.521 | 7.823 | 0.0005 |  | 7.2884 | 1.0144 | 7284.0089 | 1013.8177 |  | 29.8258 | 0.0099883 |  | 0.0040947 | 0.00056992 | 0.0040947 | 0.00078369 |  |  |  |  | 0.0046549 | -0.00056021 | 176 | #NUM!
| Bi207 | 4 |  |  |  | 52.35 | 0.5831 |  | 73.063 | 2.11 | 20.0995 | 0.0005 |  | 7.2884 | 1.0144 | 3815.4679 | 532.7268 |  | 106.3061 | 0.025329 |  | 0.027862 | 0.0038902 | 0.027862 | 0.013687 |  |  |  |  | 0.081296 | -0.053434 | 181 | #NUM!
| Bi207 | 7 |  | 569.698 | 0.002 | 97.76 | 0.03 |  | 570.223 | 1.702 | 8.266 | 0.0005 |  | 7.2884 | 1.0144 | 7125.1222 | 991.6623 |  | 35.265 | 0.010577 |  | 0.0049494 | 0.00068885 | 0.0049494 | 0.0017537 |  |  |  |  | 0.010417 | -0.0054671 | 186 | #NUM!
| Bi207 | 9 |  | 1063.656 | 0.003 | 74.58 | 0.22 |  | 1065.139 | 1.925 | 3.2037 | 0.0005 |  | 7.2884 | 1.0144 | 5435.6753 | 756.696 |  | 15.4588 | 0.0046844 |  | 0.002844 | 0.00039591 | 0.002844 | 0.00093884 |  |  |  |  | 0.0055765 | -0.0027325 | 191 | #NUM!
| Bi207 | 11 |  | 1770.228 | 0.009 | 6.871 | 0.026 |  | 1773.194 | 2.29 | 0.15869 | 0.0005 |  | 7.2884 | 1.0144 | 500.7847 | 69.7239 |  | 0.91091 | 0.002877 |  | 0.001819 | 0.00025332 | 0.001819 | 0.00056395 |  |  |  |  | 0.0033497 | -0.0015308 | 196 | #NUM!
[32 empty] | 201 | 0.2
[9 empty] | 1174.989 | 1.43 | 0.385 | [20 empty] | 206 | 0.08165
[9 empty] | 1334.42 | 1.47 | 0.339 | [20 empty] | 211 | 0.060302
|  |  |  | 1 |  |  | 2 |  |  | 3 |  |  | 4 |  |  | 5 |  |  | 6 | [12 empty] | 216 | 0.05
| X-Rays |  |  | energy/keV | branching | unc | energy/keV | branching | unc | energy/keV | branching | unc | energy/keV | branching | unc | energy/keV | branching | unc | energy/keV | branching | unc |  | weighted average energy | [8 empty] | 221 | 0.043644
[32 empty] | 226 | 0.039223
| Eu152 | 0 |  | 39.5229 | 20.8 | 0.3 | 40.1186 | 37.7 | 0.5 | [13 empty] | 39.9068 | [8 empty] | 231 | 0.035921
| Eu152 | 1 |  | 45.289 | 11.78 | 0.19 | 45.413 | 11.78 | 0.19 | 45.731 | 11.78 | 0.19 | 46.575 | 3.04 | 0.08 | 46.705 | 3.04 | 0.08 | 46.813 | 3.04 | 0.08 |  | 45.7279 | [8 empty] | 236 | 0.033333
[32 empty] | 241 | 0.031235
| Ba133 | 1 |  | 30.6254 | 33.8 | 0.4 | 30.9731 | 62.4 | 0.7 | [22 empty] | 246 | 0.029488
|  | there are others that aren't relevant to this | [29 empty] | 251 | 0.028006
[32 empty] | 256 | 0.026726
[32 empty] | 261 | 0.025607
[32 empty] | 266 | 0.024618
[32 empty] | 271 | 0.023736
[32 empty] | 276 | 0.022942
[32 empty] | 281 | 0.022222
[32 empty] | 286 | 0.021567
[32 empty] | 291 | 0.020966
[32 empty] | 296 | 0.020412
[32 empty] | 301 | 0.019901
|  | Ba-133 | ~160 | [28 empty] | 306 | 0.019426
|  | missing? | ~220 | [28 empty] | 311 | 0.018983
[32 empty] | 316 | 0.01857
[32 empty] | 321 | 0.018182
[32 empty] | 326 | 0.017817
[32 empty] | 331 | 0.017474
[32 empty] | 336 | 0.01715
[32 empty] | 341 | 0.016843
[32 empty] | 346 | 0.016552
[32 empty] | 351 | 0.016276
[32 empty] | 356 | 0.016013
[32 empty] | 361 | 0.015762
[32 empty] | 366 | 0.015523
[32 empty] | 371 | 0.015294
[32 empty] | 376 | 0.015076
[32 empty] | 381 | 0.014866
[32 empty] | 386 | 0.014665
[32 empty] | 391 | 0.014471
[32 empty] | 396 | 0.014286
[32 empty] | 401 | 0.014107
[32 empty] | 406 | 0.013935
[32 empty] | 411 | 0.013769
[32 empty] | 416 | 0.013608
[32 empty] | 421 | 0.013453
[32 empty] | 426 | 0.013304
[32 empty] | 431 | 0.013159
[32 empty] | 436 | 0.013019
[32 empty] | 441 | 0.012883
[32 empty] | 446 | 0.012752
[32 empty] | 451 | 0.012624
[32 empty] | 456 | 0.0125
[32 empty] | 461 | 0.01238
[32 empty] | 466 | 0.012263
[32 empty] | 471 | 0.012149
[32 empty] | 476 | 0.012039
[32 empty] | 481 | 0.011931
[32 empty] | 486 | 0.011826
[32 empty] | 491 | 0.011724
[32 empty] | 496 | 0.011625
[32 empty] | 501 | 0.011528
[32 empty] | 506 | 0.011433
[32 empty] | 511 | 0.011341
[32 empty] | 516 | 0.011251
[32 empty] | 521 | 0.011163
[32 empty] | 526 | 0.011077
[32 empty] | 531 | 0.010993
[32 empty] | 536 | 0.010911
[32 empty] | 541 | 0.010831
[32 empty] | 546 | 0.010752
[32 empty] | 551 | 0.010675
[32 empty] | 556 | 0.0106
[32 empty] | 561 | 0.010526
[32 empty] | 566 | 0.010454
[32 empty] | 571 | 0.010383
[32 empty] | 576 | 0.010314
[32 empty] | 581 | 0.010246
[32 empty] | 586 | 0.01018
[32 empty] | 591 | 0.010114
[32 empty] | 596 | 0.01005
[32 empty] | 601 | 0.0099875
[32 empty] | 606 | 0.0099258
[32 empty] | 611 | 0.0098653
[32 empty] | 616 | 0.0098058
[32 empty] | 621 | 0.0097474
[32 empty] | 626 | 0.00969
[32 empty] | 631 | 0.0096337
[32 empty] | 636 | 0.0095783
[32 empty] | 641 | 0.0095238
[32 empty] | 646 | 0.0094703
[32 empty] | 651 | 0.0094176
[32 empty] | 656 | 0.0093659
[32 empty] | 661 | 0.0093149
[32 empty] | 666 | 0.0092648
[32 empty] | 671 | 0.0092155
[32 empty] | 676 | 0.009167
[32 empty] | 681 | 0.0091192
[32 empty] | 686 | 0.0090722
[32 empty] | 691 | 0.0090259
[32 empty] | 696 | 0.0089803
[32 empty] | 701 | 0.0089353
[32 empty] | 706 | 0.0088911
[32 empty] | 711 | 0.0088475
[32 empty] | 716 | 0.0088045
[32 empty] | 721 | 0.0087622
[32 empty] | 726 | 0.0087204
[32 empty] | 731 | 0.0086793
[32 empty] | 736 | 0.0086387
[32 empty] | 741 | 0.0085987
[32 empty] | 746 | 0.0085592
[32 empty] | 751 | 0.0085203
[32 empty] | 756 | 0.0084819
[32 empty] | 761 | 0.008444
[32 empty] | 766 | 0.0084066
[32 empty] | 771 | 0.0083697
[32 empty] | 776 | 0.0083333
[32 empty] | 781 | 0.0082974
[32 empty] | 786 | 0.0082619
[32 empty] | 791 | 0.0082269
[32 empty] | 796 | 0.0081923
[32 empty] | 801 | 0.0081582
[32 empty] | 806 | 0.0081244
[32 empty] | 811 | 0.0080911
[32 empty] | 816 | 0.0080582
[32 empty] | 821 | 0.0080257
[32 empty] | 826 | 0.0079936
[32 empty] | 831 | 0.0079619
[32 empty] | 836 | 0.0079305
[32 empty] | 841 | 0.0078995
[32 empty] | 846 | 0.0078689
[32 empty] | 851 | 0.0078386
[32 empty] | 856 | 0.0078087
[32 empty] | 861 | 0.0077791
[32 empty] | 866 | 0.0077498
[32 empty] | 871 | 0.0077209
[32 empty] | 876 | 0.0076923
[32 empty] | 881 | 0.007664
[32 empty] | 886 | 0.007636
[32 empty] | 891 | 0.0076084
[32 empty] | 896 | 0.007581
[32 empty] | 901 | 0.0075539
[32 empty] | 906 | 0.0075271
[32 empty] | 911 | 0.0075006
[32 empty] | 916 | 0.0074744
[32 empty] | 921 | 0.0074484
[32 empty] | 926 | 0.0074227
[32 empty] | 931 | 0.0073973
[32 empty] | 936 | 0.0073721
[32 empty] | 941 | 0.0073472
[32 empty] | 946 | 0.0073225
[32 empty] | 951 | 0.0072981
[32 empty] | 956 | 0.0072739
[32 empty] | 961 | 0.00725
[32 empty] | 966 | 0.0072263
[32 empty] | 971 | 0.0072028
[32 empty] | 976 | 0.0071796
[32 empty] | 981 | 0.0071566
[32 empty] | 986 | 0.0071338
[32 empty] | 991 | 0.0071112
[32 empty] | 996 | 0.0070888
[32 empty] | 1001 | 0.0070667
[32 empty] | 1006 | 0.0070447
[32 empty] | 1011 | 0.0070229
[32 empty] | 1016 | 0.0070014
[32 empty] | 1021 | 0.00698
[32 empty] | 1026 | 0.0069589
[32 empty] | 1031 | 0.0069379
[32 empty] | 1036 | 0.0069171
[32 empty] | 1041 | 0.0068966
[32 empty] | 1046 | 0.0068761
[32 empty] | 1051 | 0.0068559
[32 empty] | 1056 | 0.0068359
[32 empty] | 1061 | 0.006816
[32 empty] | 1066 | 0.0067963
[32 empty] | 1071 | 0.0067767
[32 empty] | 1076 | 0.0067574
[32 empty] | 1081 | 0.0067382
[32 empty] | 1086 | 0.0067191
[32 empty] | 1091 | 0.0067003
[32 empty] | 1096 | 0.0066815
[32 empty] | 1101 | 0.006663
[32 empty] | 1106 | 0.0066446
[32 empty] | 1111 | 0.0066263
[32 empty] | 1116 | 0.0066082
[32 empty] | 1121 | 0.0065902
[32 empty] | 1126 | 0.0065724
[32 empty] | 1131 | 0.0065547
[32 empty] | 1136 | 0.0065372
[32 empty] | 1141 | 0.0065198
[32 empty] | 1146 | 0.0065026
[32 empty] | 1151 | 0.0064854
[32 empty] | 1156 | 0.0064685
[32 empty] | 1161 | 0.0064516
[32 empty] | 1166 | 0.0064349
[32 empty] | 1171 | 0.0064183
[32 empty] | 1176 | 0.0064018
[32 empty] | 1181 | 0.0063855
[32 empty] | 1186 | 0.0063693
[32 empty] | 1191 | 0.0063532
[32 empty] | 1196 | 0.0063372
[32 empty] | 1201 | 0.0063214
[32 empty] | 1206 | 0.0063057
[32 empty] | 1211 | 0.0062901
[32 empty] | 1216 | 0.0062746
[32 empty] | 1221 | 0.0062592
[32 empty] | 1226 | 0.0062439
[32 empty] | 1231 | 0.0062287
[32 empty] | 1236 | 0.0062137
[32 empty] | 1241 | 0.0061988
[32 empty] | 1246 | 0.0061839
[32 empty] | 1251 | 0.0061692
[32 empty] | 1256 | 0.0061546
[32 empty] | 1261 | 0.0061401
[32 empty] | 1266 | 0.0061256
[32 empty] | 1271 | 0.0061113
[32 empty] | 1276 | 0.0060971
[32 empty] | 1281 | 0.006083
[32 empty] | 1286 | 0.006069
[32 empty] | 1291 | 0.006055
[32 empty] | 1296 | 0.0060412
[32 empty] | 1301 | 0.0060275
[32 empty] | 1306 | 0.0060138
[32 empty] | 1311 | 0.0060003
[32 empty] | 1316 | 0.0059868
[32 empty] | 1321 | 0.0059735
[32 empty] | 1326 | 0.0059602
[32 empty] | 1331 | 0.005947
[32 empty] | 1336 | 0.0059339
[32 empty] | 1341 | 0.0059209
[32 empty] | 1346 | 0.005908
[32 empty] | 1351 | 0.0058951
[32 empty] | 1356 | 0.0058824
[32 empty] | 1361 | 0.0058697
[32 empty] | 1366 | 0.0058571
[32 empty] | 1371 | 0.0058446
[32 empty] | 1376 | 0.0058321
[32 empty] | 1381 | 0.0058198
[32 empty] | 1386 | 0.0058075
[32 empty] | 1391 | 0.0057953
[32 empty] | 1396 | 0.0057831
[32 empty] | 1401 | 0.0057711
[32 empty] | 1406 | 0.0057591
[32 empty] | 1411 | 0.0057472
[32 empty] | 1416 | 0.0057354
[32 empty] | 1421 | 0.0057236
[32 empty] | 1426 | 0.005712
[32 empty] | 1431 | 0.0057003
[32 empty] | 1436 | 0.0056888
[32 empty] | 1441 | 0.0056773
[32 empty] | 1446 | 0.0056659
[32 empty] | 1451 | 0.0056546
[32 empty] | 1456 | 0.0056433
[32 empty] | 1461 | 0.0056321
[32 empty] | 1466 | 0.005621
[32 empty] | 1471 | 0.0056099
[32 empty] | 1476 | 0.0055989
[32 empty] | 1481 | 0.005588
[32 empty] | 1486 | 0.0055771
[32 empty] | 1491 | 0.0055663
[32 empty] | 1496 | 0.0055556
[32 empty] | 1501 | 0.0055449
[32 empty] | 1506 | 0.0055342
[32 empty] | 1511 | 0.0055237
[32 empty] | 1516 | 0.0055132
[32 empty] | 1521 | 0.0055027
[32 empty] | 1526 | 0.0054924
[32 empty] | 1531 | 0.005482
[32 empty] | 1536 | 0.0054718
[32 empty] | 1541 | 0.0054615
[32 empty] | 1546 | 0.0054514
[32 empty] | 1551 | 0.0054413
[32 empty] | 1556 | 0.0054313
[32 empty] | 1561 | 0.0054213
[32 empty] | 1566 | 0.0054113
[32 empty] | 1571 | 0.0054015
[32 empty] | 1576 | 0.0053916
[32 empty] | 1581 | 0.0053819
[32 empty] | 1586 | 0.0053722
[32 empty] | 1591 | 0.0053625
[32 empty] | 1596 | 0.0053529
[32 empty] | 1601 | 0.0053433
[32 empty] | 1606 | 0.0053338
[32 empty] | 1611 | 0.0053243
[32 empty] | 1616 | 0.0053149
[32 empty] | 1621 | 0.0053056
[32 empty] | 1626 | 0.0052963
[32 empty] | 1631 | 0.005287
[32 empty] | 1636 | 0.0052778
[32 empty] | 1641 | 0.0052686
[32 empty] | 1646 | 0.0052595
[32 empty] | 1651 | 0.0052504
[32 empty] | 1656 | 0.0052414
[32 empty] | 1661 | 0.0052324
[32 empty] | 1666 | 0.0052235
[32 empty] | 1671 | 0.0052146
[32 empty] | 1676 | 0.0052058
[32 empty] | 1681 | 0.005197
[32 empty] | 1686 | 0.0051882
[32 empty] | 1691 | 0.0051795
[32 empty] | 1696 | 0.0051709
[32 empty] | 1701 | 0.0051623
[32 empty] | 1706 | 0.0051537
[32 empty] | 1711 | 0.0051451
[32 empty] | 1716 | 0.0051367
[32 empty] | 1721 | 0.0051282
[32 empty] | 1726 | 0.0051198
[32 empty] | 1731 | 0.0051114
[32 empty] | 1736 | 0.0051031
[32 empty] | 1741 | 0.0050948
[32 empty] | 1746 | 0.0050866
[32 empty] | 1751 | 0.0050784
[32 empty] | 1756 | 0.0050702
[32 empty] | 1761 | 0.0050621
[32 empty] | 1766 | 0.005054
[32 empty] | 1771 | 0.0050459
[32 empty] | 1776 | 0.0050379
[32 empty] | 1781 | 0.00503
[32 empty] | 1786 | 0.005022
[32 empty] | 1791 | 0.0050141
[32 empty] | 1796 | 0.0050063
[32 empty] | 1801 | 0.0049984
[32 empty] | 1806 | 0.0049907
[32 empty] | 1811 | 0.0049829
[32 empty] | 1816 | 0.0049752
[32 empty] | 1821 | 0.0049675
[32 empty] | 1826 | 0.0049599
[32 empty] | 1831 | 0.0049523
[32 empty] | 1836 | 0.0049447
[32 empty] | 1841 | 0.0049371
[32 empty] | 1846 | 0.0049296
[32 empty] | 1851 | 0.0049222
[32 empty] | 1856 | 0.0049147
[32 empty] | 1861 | 0.0049073
[32 empty] | 1866 | 0.0049
[32 empty] | 1871 | 0.0048926
[32 empty] | 1876 | 0.0048853
[32 empty] | 1881 | 0.004878
[32 empty] | 1886 | 0.0048708
[32 empty] | 1891 | 0.0048636
[32 empty] | 1896 | 0.0048564
[32 empty] | 1901 | 0.0048493
[32 empty] | 1906 | 0.0048422
[32 empty] | 1911 | 0.0048351
[32 empty] | 1916 | 0.004828
[32 empty] | 1921 | 0.004821
[32 empty] | 1926 | 0.004814
[32 empty] | 1931 | 0.0048071
[32 empty] | 1936 | 0.0048002
[32 empty] | 1941 | 0.0047933
[32 empty] | 1946 | 0.0047864
[32 empty] | 1951 | 0.0047795
[32 empty] | 1956 | 0.0047727
[32 empty] | 1961 | 0.004766
[32 empty] | 1966 | 0.0047592

## Energy Efficiency Calibrati (2)
time/s | isotope | peak |  | literature energy /keV | uncertainty | photons per 100 disintegrations | uncertainty |  | mean /keV | standard deviation | unc | amplitude | unc |  | current activity /kBq | uncertainty | expected area | uncertainty |  | measured area | unc |  | ratio | unc | y | 1/x | [5 empty] | fit | residual | fit line x | fit line y
[30 empty] | max | min
1943 | Ba133 | 1 |  |  |  | 96.2 | 0.80623 |  | 30.983 | 1.197 | 0.055159 | 160.4152 | 0.14367 |  | 9.403 | 0.58338 | 9045.7337 | 566.3086 |  | 481.3153 | 22.1838 |  | 0.053209 | 0.0041365 | 0.053209 | 0.032276 | a | 6.7 | 0.2 | 6.5 | 6.9 | 0.073138 | -0.019929 | 1 | -5993.3
1943 | Ba133 | 5 |  | 80.9979 | 0.0011 | 33.31 | 0.3 |  | 81.23 | 1.452 | 0.06691 | 56.1366 | 0.084988 |  | 9.403 | 0.58338 | 3132.1558 | 196.3607 |  | 204.3162 | 9.4202 |  | 0.065232 | 0.0050764 | 0.065232 | 0.012311 | b | 3.1 | 0.1 | 3.2 | 3 | 0.07527 | -0.010038 | 6 | -22.1044
1943 | Ba133 | 7 |  | 160.6121 | 0.0016 | 0.638 | 0.006 |  | 161.292 | 1.233 | 0.056818 | 1.2298 | 0.012579 |  | 9.403 | 0.58338 | 59.9915 | 3.7645 |  | 3.8008 | 0.17941 |  | 0.063356 | 0.0049748 | 0.063356 | 0.0061999 | c | 6000 | 500 | 6500 | 5500 | 0.040679 | 0.022677 | 11 | -2.9377
1943 | Ba133 | 9 |  | 276.3989 | 0.0012 | 7.13 | 0.06 |  | 277.546 | 1.201 | 0.055343 | 6.0671 | 0.02794 |  | 9.403 | 0.58338 | 670.4374 | 41.9759 |  | 18.2648 | 0.84586 |  | 0.027243 | 0.0021216 | 0.027243 | 0.003603 | d | 1 |  |  |  | 0.02398 | 0.0032629 | 16 | -0.69139
1943 | Ba133 | 10 |  | 302.8508 | 0.0005 | 18.31 | 0.11 |  | 304.1 | 1.198 | 0.055205 | 14.8318 | 0.043685 |  | 9.403 | 0.58338 | 1721.6984 | 107.3165 |  | 44.5391 | 2.0566 |  | 0.025869 | 0.0020067 | 0.025869 | 0.0032884 | [5 empty] | 0.021912 | 0.0039575 | 21 | -0.15878
1943 | Ba133 | 12 |  | 356.0129 | 0.0007 | 62.05 | 0.19 |  | 357.4 | 1.45 | 0.066818 | 33.1124 | 0.065272 |  | 9.403 | 0.58338 | 5834.5923 | 362.4278 |  | 120.3508 | 5.551 |  | 0.020627 | 0.0015959 | 0.020627 | 0.002798 | y = a/(x^d) - c/(x^b) |  |  |  |  | 0.018673 | 0.0019536 | 26 | 0.011239
1943 | Ba133 | 13 |  | 383.8485 | 0.0012 | 8.94 | 0.06 |  | 385.95 | 1.86 | 0.085711 | 6.0137 | 0.027817 |  | 9.403 | 0.58338 | 840.6326 | 52.4584 |  | 28.0378 | 1.2985 |  | 0.033353 | 0.0025919 | 0.033353 | 0.002591 | [5 empty] | 0.017302 | 0.016051 | 31 | 0.073263
7876 | Eu152 | 1 |  |  |  | 58.5 | 0.5831 |  | 39.907 | 1.181 | 0.054422 | 48.9149 | 0.039404 |  | 5.2714 | 0.15293 | 3083.753 | 94.5988 |  | 144.8043 | 6.6738 |  | 0.046957 | 0.0025997 | 0.046957 | 0.025058 | [5 empty] | 0.10259 | -0.055635 | 36 | 0.096241
2625 | Co60 | 1 |  | 1173.228 | 0.003 | 99.85 | 0.03 |  | 1176.789 | 1.648 | 0.075942 | 19.3525 | 0.042931 |  | 10.0749 | 0.72812 | 10059.8292 | 727.033 |  | 79.9438 | 3.6882 |  | 0.0079468 | 0.00068137 | 0.0079468 | 0.00085107 | [5 empty] | 0.0056916 | 0.0022552 | 41 | 0.10336
2625 | Co60 | 2 |  | 1332.492 | 0.004 | 99.9826 | 0.0006 |  | 1336.513 | 1.704 | 0.078522 | 16.7377 | 0.039926 |  | 10.0749 | 0.72812 | 10073.1886 | 727.9922 |  | 71.4915 | 3.2988 |  | 0.0070972 | 0.00060855 | 0.0070972 | 0.00074939 | [5 empty] | 0.0050118 | 0.0020854 | 46 | 0.10362
6516 | Cs137 | 1 |  | 661.657 | 0.003 | 84.99 | 0.2 |  | 662.747 | 1.26 | 0.058062 | 50.4827 | 0.04401 |  | 15.2714 | 0.93406 | 12979.1607 | 794.4413 |  | 159.4421 | 7.3486 |  | 0.012284 | 0.00094125 | 0.012284 | 0.0015089 | [5 empty] | 0.010099 | 0.0021858 | 51 | 0.10085
7876 | Eu152 | 7 |  | 121.7817 | 0.0003 | 28.41 | 0.13 |  | 122.207 | 1.128 | 0.05198 | 23.4152 | 0.027263 |  | 5.2714 | 0.15293 | 1497.597 | 43.9855 |  | 66.2058 | 3.0518 |  | 0.044208 | 0.0024163 | 0.044208 | 0.0081828 | [5 empty] | 0.052792 | -0.0085839 | 56 | 0.096799
7876 | Eu152 | 9 |  | 244.6974 | 0.0008 | 7.55 | 0.04 |  | 245.391 | 1.161 | 0.0535 | 3.1666 | 0.010026 |  | 5.2714 | 0.15293 | 397.9886 | 11.7374 |  | 9.2154 | 0.42566 |  | 0.023155 | 0.0012689 | 0.023155 | 0.0040751 | [5 empty] | 0.027069 | -0.0039142 | 61 | 0.092312
7876 | Eu152 | 12 |  | 344.2785 | 0.0012 | 26.59 | 0.12 |  | 345.095 | 1.182 | 0.054468 | 8.7809 | 0.016695 |  | 5.2714 | 0.15293 | 1401.658 | 41.154 |  | 26.0165 | 1.1999 |  | 0.018561 | 0.0010148 | 0.018561 | 0.0028978 | [5 empty] | 0.019334 | -0.00077236 | 66 | 0.087788
7876 | Eu152 | 13 |  | 367.7891 | 0.002 | 0.862 | 0.005 |  | 368.299 | 1.41 | 0.064974 | 0.29648 | 0.0030677 |  | 5.2714 | 0.15293 | 45.4392 | 1.3444 |  | 1.0479 | 0.04949 |  | 0.023061 | 0.0012852 | 0.023061 | 0.0027152 | [5 empty] | 0.018125 | 0.0049358 | 71 | 0.08342
7876 | Eu152 | 14 |  | 411.1165 | 0.0012 | 2.238 | 0.01 |  | 412.016 | 1.223 | 0.056357 | 0.55991 | 0.0042158 |  | 5.2714 | 0.15293 | 117.9733 | 3.463 |  | 1.7165 | 0.080146 |  | 0.01455 | 0.00080245 | 0.01455 | 0.0024271 | [5 empty] | 0.016215 | -0.0016649 | 76 | 0.079294
7876 | Eu152 | 15 |  | 443.965 | 0.003 | 2.8 | 0.02 |  | 444.873 | 1.223 | 0.056357 | 0.68236 | 0.004654 |  | 5.2714 | 0.15293 | 147.5984 | 4.41 |  | 2.0918 | 0.097445 |  | 0.014173 | 0.00078433 | 0.014173 | 0.0022478 | [5 empty] | 0.015023 | -0.00085094 | 81 | 0.075441
7876 | Eu152 | 16 |  | 778.9045 | 0.0024 | 12.97 | 0.06 |  | 780.077 | 1.31 | 0.060366 | 1.6744 | 0.0072903 |  | 5.2714 | 0.15293 | 683.697 | 20.086 |  | 5.4981 | 0.25449 |  | 0.0080417 | 0.00044087 | 0.0080417 | 0.0012819 | [5 empty] | 0.0085824 | -0.00054067 | 86 | 0.071865
7876 | Eu152 | 17 |  | 867.38 | 0.003 | 4.243 | 0.023 |  | 868.651 | 1.338 | 0.061657 | 0.42262 | 0.0036626 |  | 5.2714 | 0.15293 | 223.6643 | 6.6013 |  | 1.4174 | 0.066461 |  | 0.0063372 | 0.00035111 | 0.0063372 | 0.0011512 | [5 empty] | 0.0077085 | -0.0013712 | 91 | 0.068555
7876 | Eu152 | 18 |  | 964.079 | 0.018 | 14.5 | 0.06 |  | 965.354 | 1.423 | 0.065574 | 1.3572 | 0.0065636 |  | 5.2714 | 0.15293 | 764.349 | 22.3998 |  | 4.8411 | 0.22431 |  | 0.0063336 | 0.00034723 | 0.0063336 | 0.0010359 | [5 empty] | 0.0069371 | -0.00060349 | 96 | 0.065495
7876 | Eu152 | 19 |  | 1085.837 | 0.01 | 10.13 | 0.06 |  | 1087.38 | 1.605 | 0.07396 | 0.88351 | 0.0052957 |  | 5.2714 | 0.15293 | 533.99 | 15.8117 |  | 3.5545 | 0.16518 |  | 0.0066565 | 0.00036678 | 0.0066565 | 0.00091964 | [5 empty] | 0.0061593 | 0.00049722 | 101 | 0.062666
7876 | Eu152 | 20 |  | 1112.076 | 0.003 | 13.41 | 0.06 |  | 1113.481 | 1.445 | 0.066587 | 1.1038 | 0.0059191 |  | 5.2714 | 0.15293 | 706.8911 | 20.7508 |  | 3.998 | 0.18548 |  | 0.0056557 | 0.0003105 | 0.0056557 | 0.00089808 | [5 empty] | 0.006015 | -0.00035927 | 106 | 0.060047
7876 | Eu152 | 21 |  | 1408.013 | 0.003 | 20.85 | 0.08 |  | 1409.718 | 1.59 | 0.073269 | 1.2346 | 0.0062601 |  | 5.2714 | 0.15293 | 1099.0812 | 32.1643 |  | 4.9206 | 0.22812 |  | 0.004477 | 0.00024544 | 0.004477 | 0.00070936 | [5 empty] | 0.0047517 | -0.00027468 | 111 | 0.057621
610 | Na22 | 1 |  | 511 | 0 | 180.7 | 0.2 |  | 511.869 | 1.765 | 0.081333 | 26.9006 | 0.105 |  | 7.2884 | 1.0144 | 13170.1062 | 1833.0463 |  | 119.0138 | 5.5039 |  | 0.0090367 | 0.0013254 | 0.0090367 | 0.0019536 | [5 empty] | 0.013065 | -0.0040287 | 116 | 0.055369
610 | Na22 | 2 |  | 1274.537 | 0.007 | 99.94 | 0.13 |  | 1276.021 | 1.524 | 0.070228 | 7.8349 | 0.056666 |  | 7.2884 | 1.0144 | 7284.0089 | 1013.8177 |  | 29.9302 | 1.3961 |  | 0.004109 | 0.00060317 | 0.004109 | 0.00078369 | [5 empty] | 0.0052493 | -0.0011403 | 121 | 0.053275
2462 | Bi207 | 4 |  |  |  | 52.35 | 0.5831 |  | 72.971 | 2.462 | 0.11345 | 175.3889 | 0.13345 |  | 49.2371 | 0.22754 | 25775.6272 | 310.8291 |  | 1082.3805 | 49.8842 |  | 0.041992 | 0.0020005 | 0.041992 | 0.013704 | [5 empty] | 0.081762 | -0.03977 | 126 | 0.051325
2462 | Bi207 | 7 |  | 569.698 | 0.002 | 97.76 | 0.03 |  | 570.175 | 1.891 | 0.087139 | 78.8147 | 0.08946 |  | 49.2371 | 0.22754 | 48134.199 | 222.9319 |  | 373.5843 | 17.2204 |  | 0.0077613 | 0.00035956 | 0.0077613 | 0.0017538 | [5 empty] | 0.011734 | -0.0039723 | 131 | 0.049506
2462 | Bi207 | 9 |  | 1063.656 | 0.003 | 74.58 | 0.22 |  | 1065.09 | 2.132 | 0.098245 | 30.631 | 0.055771 |  | 49.2371 | 0.22754 | 36721.0368 | 201.3235 |  | 163.6961 | 7.5492 |  | 0.0044578 | 0.00020703 | 0.0044578 | 0.00093889 | [5 empty] | 0.0062881 | -0.0018302 | 136 | 0.047805
2462 | Bi207 | 11 |  | 1770.228 | 0.009 | 6.871 | 0.026 |  | 1773.217 | 2.183 | 0.1006 | 1.7886 | 0.013477 |  | 49.2371 | 0.22754 | 3383.0818 | 20.2067 |  | 9.7871 | 0.45699 |  | 0.002893 | 0.00013618 | 0.002893 | 0.00056395 | [5 empty] | 0.0037779 | -0.00088498 | 141 | 0.046213
1390 | Am241 | 1 |  | 59.5409 | 0.0001 | 35.92 | 0.17 |  | 59.683 | 1.309 | 0.06032 | 432.1834 | 0.2788 |  | 36.648 | 1.4662 | 13163.9628 | 530.3368 |  | 1418.0699 | 65.3527 |  | 0.10772 | 0.006594 | 0.10772 | 0.016755 | [5 empty] | 0.093509 | 0.014214 | 146 | 0.044719
[34 empty] | 151 | 0.043316
[9 empty] | 1174.989 | 1.43 |  | 0.385 | [21 empty] | 156 | 0.041995
[9 empty] | 1334.42 | 1.47 |  | 0.339 | [21 empty] | 161 | 0.04075
|  |  |  | 1 |  |  | 2 |  |  | 3 |  |  |  | 4 |  |  | 5 |  |  | 6 | [13 empty] | 166 | 0.039575
| X-Rays |  |  | energy/keV | branching | unc | energy/keV | branching | unc | energy/keV |  | branching | unc | energy/keV | branching | unc | energy/keV | branching | unc | energy/keV | branching | unc |  | weighted average energy | [9 empty] | 171 | 0.038464
[34 empty] | 176 | 0.037412
| Eu152 | 0 |  | 39.5229 | 20.8 | 0.3 | 40.1186 | 37.7 | 0.5 | [14 empty] | 39.9068 | [9 empty] | 181 | 0.036415
| Eu152 | 1 |  | 45.289 | 11.78 | 0.19 | 45.413 | 11.78 | 0.19 | 45.731 |  | 11.78 | 0.19 | 46.575 | 3.04 | 0.08 | 46.705 | 3.04 | 0.08 | 46.813 | 3.04 | 0.08 |  | 45.7279 | [9 empty] | 186 | 0.035469
[34 empty] | 191 | 0.034569
| Ba133 | 1 |  | 30.6254 | 33.8 | 0.4 | 30.9731 | 62.4 | 0.7 | [24 empty] | 196 | 0.033714
|  | there are others that aren't relevant to this | [31 empty] | 201 | 0.032899
[34 empty] | 206 | 0.032121
[34 empty] | 211 | 0.03138
[34 empty] | 216 | 0.030671
[34 empty] | 221 | 0.029993
[34 empty] | 226 | 0.029344
[34 empty] | 231 | 0.028722
[34 empty] | 236 | 0.028126
[34 empty] | 241 | 0.027553
[34 empty] | 246 | 0.027003
[34 empty] | 251 | 0.026475
|  | Ba-133 | ~160 | [30 empty] | 256 | 0.025966
|  | missing? | ~220 | [30 empty] | 261 | 0.025477
[34 empty] | 266 | 0.025006
[34 empty] | 271 | 0.024551
[34 empty] | 276 | 0.024113
[34 empty] | 281 | 0.02369
[34 empty] | 286 | 0.023281
[34 empty] | 291 | 0.022886
[34 empty] | 296 | 0.022504
[34 empty] | 301 | 0.022135
[34 empty] | 306 | 0.021777
[34 empty] | 311 | 0.021431
[34 empty] | 316 | 0.021096
[34 empty] | 321 | 0.02077
[34 empty] | 326 | 0.020455
[34 empty] | 331 | 0.020149
[34 empty] | 336 | 0.019852
[34 empty] | 341 | 0.019564
[34 empty] | 346 | 0.019283
[34 empty] | 351 | 0.019011
[34 empty] | 356 | 0.018746
[34 empty] | 361 | 0.018489
[34 empty] | 366 | 0.018238
[34 empty] | 371 | 0.017994
[34 empty] | 376 | 0.017757
[34 empty] | 381 | 0.017525
[34 empty] | 386 | 0.0173
[34 empty] | 391 | 0.01708
[34 empty] | 396 | 0.016866
[34 empty] | 401 | 0.016657
[34 empty] | 406 | 0.016453
[34 empty] | 411 | 0.016254
[34 empty] | 416 | 0.01606
[34 empty] | 421 | 0.015871
[34 empty] | 426 | 0.015685
[34 empty] | 431 | 0.015504
[34 empty] | 436 | 0.015328
[34 empty] | 441 | 0.015155
[34 empty] | 446 | 0.014986
[34 empty] | 451 | 0.01482
[34 empty] | 456 | 0.014659
[34 empty] | 461 | 0.0145
[34 empty] | 466 | 0.014346
[34 empty] | 471 | 0.014194
[34 empty] | 476 | 0.014046
[34 empty] | 481 | 0.0139
[34 empty] | 486 | 0.013758
[34 empty] | 491 | 0.013618
[34 empty] | 496 | 0.013482
[34 empty] | 501 | 0.013348
[34 empty] | 506 | 0.013216
[34 empty] | 511 | 0.013087
[34 empty] | 516 | 0.012961
[34 empty] | 521 | 0.012837
[34 empty] | 526 | 0.012716
[34 empty] | 531 | 0.012596
[34 empty] | 536 | 0.012479
[34 empty] | 541 | 0.012364
[34 empty] | 546 | 0.012251
[34 empty] | 551 | 0.012141
[34 empty] | 556 | 0.012032
[34 empty] | 561 | 0.011925
[34 empty] | 566 | 0.01182
[34 empty] | 571 | 0.011717
[34 empty] | 576 | 0.011615
[34 empty] | 581 | 0.011516
[34 empty] | 586 | 0.011418
[34 empty] | 591 | 0.011321
[34 empty] | 596 | 0.011227
[34 empty] | 601 | 0.011134
[34 empty] | 606 | 0.011042
[34 empty] | 611 | 0.010952
[34 empty] | 616 | 0.010863
[34 empty] | 621 | 0.010776
[34 empty] | 626 | 0.01069
[34 empty] | 631 | 0.010606
[34 empty] | 636 | 0.010522
[34 empty] | 641 | 0.01044
[34 empty] | 646 | 0.01036
[34 empty] | 651 | 0.01028
[34 empty] | 656 | 0.010202
[34 empty] | 661 | 0.010125
[34 empty] | 666 | 0.010049
[34 empty] | 671 | 0.0099747
[34 empty] | 676 | 0.0099011
[34 empty] | 681 | 0.0098286
[34 empty] | 686 | 0.0097571
[34 empty] | 691 | 0.0096866
[34 empty] | 696 | 0.0096172
[34 empty] | 701 | 0.0095487
[34 empty] | 706 | 0.0094812
[34 empty] | 711 | 0.0094147
[34 empty] | 716 | 0.0093491
[34 empty] | 721 | 0.0092844
[34 empty] | 726 | 0.0092205
[34 empty] | 731 | 0.0091576
[34 empty] | 736 | 0.0090955
[34 empty] | 741 | 0.0090342
[34 empty] | 746 | 0.0089738
[34 empty] | 751 | 0.0089141
[34 empty] | 756 | 0.0088553
[34 empty] | 761 | 0.0087972
[34 empty] | 766 | 0.0087399
[34 empty] | 771 | 0.0086833
[34 empty] | 776 | 0.0086274
[34 empty] | 781 | 0.0085723
[34 empty] | 786 | 0.0085178
[34 empty] | 791 | 0.0084641
[34 empty] | 796 | 0.008411
[34 empty] | 801 | 0.0083586
[34 empty] | 806 | 0.0083068
[34 empty] | 811 | 0.0082556
[34 empty] | 816 | 0.0082051
[34 empty] | 821 | 0.0081552
[34 empty] | 826 | 0.0081059
[34 empty] | 831 | 0.0080572
[34 empty] | 836 | 0.0080091
[34 empty] | 841 | 0.0079616
[34 empty] | 846 | 0.0079146
[34 empty] | 851 | 0.0078681
[34 empty] | 856 | 0.0078222
[34 empty] | 861 | 0.0077769
[34 empty] | 866 | 0.007732
[34 empty] | 871 | 0.0076877
[34 empty] | 876 | 0.0076439
[34 empty] | 881 | 0.0076005
[34 empty] | 886 | 0.0075577
[34 empty] | 891 | 0.0075153
[34 empty] | 896 | 0.0074735
[34 empty] | 901 | 0.007432
[34 empty] | 906 | 0.0073911
[34 empty] | 911 | 0.0073505
[34 empty] | 916 | 0.0073105
[34 empty] | 921 | 0.0072708
[34 empty] | 926 | 0.0072316
[34 empty] | 931 | 0.0071928
[34 empty] | 936 | 0.0071544
[34 empty] | 941 | 0.0071165
[34 empty] | 946 | 0.0070789
[34 empty] | 951 | 0.0070417
[34 empty] | 956 | 0.0070049
[34 empty] | 961 | 0.0069685
[34 empty] | 966 | 0.0069325
[34 empty] | 971 | 0.0068968
[34 empty] | 976 | 0.0068615
[34 empty] | 981 | 0.0068266
[34 empty] | 986 | 0.006792
[34 empty] | 991 | 0.0067578
[34 empty] | 996 | 0.0067239
[34 empty] | 1001 | 0.0066903
[34 empty] | 1006 | 0.0066571
[34 empty] | 1011 | 0.0066242
[34 empty] | 1016 | 0.0065916
[34 empty] | 1021 | 0.0065594
[34 empty] | 1026 | 0.0065274
[34 empty] | 1031 | 0.0064958
[34 empty] | 1036 | 0.0064645
[34 empty] | 1041 | 0.0064335
[34 empty] | 1046 | 0.0064027
[34 empty] | 1051 | 0.0063723
[34 empty] | 1056 | 0.0063422
[34 empty] | 1061 | 0.0063123
[34 empty] | 1066 | 0.0062827
[34 empty] | 1071 | 0.0062534
[34 empty] | 1076 | 0.0062244
[34 empty] | 1081 | 0.0061956
[34 empty] | 1086 | 0.0061671
[34 empty] | 1091 | 0.0061389
[34 empty] | 1096 | 0.0061109
[34 empty] | 1101 | 0.0060831
[34 empty] | 1106 | 0.0060557
[34 empty] | 1111 | 0.0060284
[34 empty] | 1116 | 0.0060014
[34 empty] | 1121 | 0.0059747
[34 empty] | 1126 | 0.0059482
[34 empty] | 1131 | 0.0059219
[34 empty] | 1136 | 0.0058959
[34 empty] | 1141 | 0.00587
[34 empty] | 1146 | 0.0058445
[34 empty] | 1151 | 0.0058191
[34 empty] | 1156 | 0.0057939
[34 empty] | 1161 | 0.005769
[34 empty] | 1166 | 0.0057443
[34 empty] | 1171 | 0.0057198
[34 empty] | 1176 | 0.0056955
[34 empty] | 1181 | 0.0056714
[34 empty] | 1186 | 0.0056475
[34 empty] | 1191 | 0.0056238
[34 empty] | 1196 | 0.0056003
[34 empty] | 1201 | 0.005577
[34 empty] | 1206 | 0.0055539
[34 empty] | 1211 | 0.005531
[34 empty] | 1216 | 0.0055082
[34 empty] | 1221 | 0.0054857
[34 empty] | 1226 | 0.0054633
[34 empty] | 1231 | 0.0054412
[34 empty] | 1236 | 0.0054192
[34 empty] | 1241 | 0.0053973
[34 empty] | 1246 | 0.0053757
[34 empty] | 1251 | 0.0053542
[34 empty] | 1256 | 0.0053329
[34 empty] | 1261 | 0.0053118
[34 empty] | 1266 | 0.0052908
[34 empty] | 1271 | 0.00527
[34 empty] | 1276 | 0.0052494
[34 empty] | 1281 | 0.0052289
[34 empty] | 1286 | 0.0052086
[34 empty] | 1291 | 0.0051884
[34 empty] | 1296 | 0.0051684
[34 empty] | 1301 | 0.0051486
[34 empty] | 1306 | 0.0051289
[34 empty] | 1311 | 0.0051093
[34 empty] | 1316 | 0.0050899
[34 empty] | 1321 | 0.0050706
[34 empty] | 1326 | 0.0050515
[34 empty] | 1331 | 0.0050326
[34 empty] | 1336 | 0.0050137
[34 empty] | 1341 | 0.0049951
[34 empty] | 1346 | 0.0049765
[34 empty] | 1351 | 0.0049581
[34 empty] | 1356 | 0.0049398
[34 empty] | 1361 | 0.0049217
[34 empty] | 1366 | 0.0049037
[34 empty] | 1371 | 0.0048858
[34 empty] | 1376 | 0.0048681
[34 empty] | 1381 | 0.0048505
[34 empty] | 1386 | 0.004833
[34 empty] | 1391 | 0.0048156
[34 empty] | 1396 | 0.0047984
[34 empty] | 1401 | 0.0047812
[34 empty] | 1406 | 0.0047642
[34 empty] | 1411 | 0.0047474
[34 empty] | 1416 | 0.0047306
[34 empty] | 1421 | 0.004714
[34 empty] | 1426 | 0.0046975
[34 empty] | 1431 | 0.0046811
[34 empty] | 1436 | 0.0046648
[34 empty] | 1441 | 0.0046486
[34 empty] | 1446 | 0.0046325
[34 empty] | 1451 | 0.0046166
[34 empty] | 1456 | 0.0046007
[34 empty] | 1461 | 0.004585
[34 empty] | 1466 | 0.0045693
[34 empty] | 1471 | 0.0045538
[34 empty] | 1476 | 0.0045384
[34 empty] | 1481 | 0.0045231
[34 empty] | 1486 | 0.0045079
[34 empty] | 1491 | 0.0044928
[34 empty] | 1496 | 0.0044777
[34 empty] | 1501 | 0.0044628
[34 empty] | 1506 | 0.004448
[34 empty] | 1511 | 0.0044333
[34 empty] | 1516 | 0.0044187
[34 empty] | 1521 | 0.0044042
[34 empty] | 1526 | 0.0043898
[34 empty] | 1531 | 0.0043754
[34 empty] | 1536 | 0.0043612
[34 empty] | 1541 | 0.004347
[34 empty] | 1546 | 0.004333
[34 empty] | 1551 | 0.004319
[34 empty] | 1556 | 0.0043051
[34 empty] | 1561 | 0.0042914
[34 empty] | 1566 | 0.0042777
[34 empty] | 1571 | 0.0042641
[34 empty] | 1576 | 0.0042505
[34 empty] | 1581 | 0.0042371
[34 empty] | 1586 | 0.0042237
[34 empty] | 1591 | 0.0042105
[34 empty] | 1596 | 0.0041973
[34 empty] | 1601 | 0.0041842
[34 empty] | 1606 | 0.0041712
[34 empty] | 1611 | 0.0041582
[34 empty] | 1616 | 0.0041454
[34 empty] | 1621 | 0.0041326
[34 empty] | 1626 | 0.0041199
[34 empty] | 1631 | 0.0041072
[34 empty] | 1636 | 0.0040947
[34 empty] | 1641 | 0.0040822
[34 empty] | 1646 | 0.0040698
[34 empty] | 1651 | 0.0040575
[34 empty] | 1656 | 0.0040453
[34 empty] | 1661 | 0.0040331
[34 empty] | 1666 | 0.004021
[34 empty] | 1671 | 0.004009
[34 empty] | 1676 | 0.003997
[34 empty] | 1681 | 0.0039851
[34 empty] | 1686 | 0.0039733
[34 empty] | 1691 | 0.0039616
[34 empty] | 1696 | 0.0039499
[34 empty] | 1701 | 0.0039383
[34 empty] | 1706 | 0.0039267
[34 empty] | 1711 | 0.0039153
[34 empty] | 1716 | 0.0039039
[34 empty] | 1721 | 0.0038925
[34 empty] | 1726 | 0.0038813
[34 empty] | 1731 | 0.00387
[34 empty] | 1736 | 0.0038589
[34 empty] | 1741 | 0.0038478
[34 empty] | 1746 | 0.0038368
[34 empty] | 1751 | 0.0038259
[34 empty] | 1756 | 0.003815
[34 empty] | 1761 | 0.0038041
[34 empty] | 1766 | 0.0037934
[34 empty] | 1771 | 0.0037827
[34 empty] | 1776 | 0.003772
[34 empty] | 1781 | 0.0037614
[34 empty] | 1786 | 0.0037509
[34 empty] | 1791 | 0.0037404
[34 empty] | 1796 | 0.00373
[34 empty] | 1801 | 0.0037197
[34 empty] | 1806 | 0.0037094
[34 empty] | 1811 | 0.0036991
[34 empty] | 1816 | 0.003689
[34 empty] | 1821 | 0.0036788
[34 empty] | 1826 | 0.0036688
[34 empty] | 1831 | 0.0036587
[34 empty] | 1836 | 0.0036488
[34 empty] | 1841 | 0.0036389
[34 empty] | 1846 | 0.003629
[34 empty] | 1851 | 0.0036192
[34 empty] | 1856 | 0.0036095
[34 empty] | 1861 | 0.0035998
[34 empty] | 1866 | 0.0035901
[34 empty] | 1871 | 0.0035805
[34 empty] | 1876 | 0.003571
[34 empty] | 1881 | 0.0035615
[34 empty] | 1886 | 0.0035521
[34 empty] | 1891 | 0.0035427
[34 empty] | 1896 | 0.0035333
[34 empty] | 1901 | 0.0035241
[34 empty] | 1906 | 0.0035148
[34 empty] | 1911 | 0.0035056
[34 empty] | 1916 | 0.0034965
[34 empty] | 1921 | 0.0034874
[34 empty] | 1926 | 0.0034783
[34 empty] | 1931 | 0.0034693
[34 empty] | 1936 | 0.0034604
[34 empty] | 1941 | 0.0034514
[34 empty] | 1946 | 0.0034426
[34 empty] | 1951 | 0.0034338
[34 empty] | 1956 | 0.003425
[34 empty] | 1961 | 0.0034163
[34 empty] | 1966 | 0.0034076

## Geometry Efficiency Calibration
[48 empty] | z = 0.7 | h1 = 1.8 | h2 = 0.7
| z/mm | r/mm | mean | std dev | unc | amplitude | unc | mean | std dev | unc | amplitude | unc | area 1 | unc | area 2 | unc | norm 1 | unc | norm 2 | unc | norm av | unc | [7 empty] | detector r/mm | 40 | 39 | 44 | 50 | 65 | 75 | fits |  | param | 2250 | 2 | [5 empty] | rho/inch |  | disk | cylinder h1 | cylinder h2 | total
[24 empty] | 0 | 10 | 20 | 30 | 40 | 50 | [15 empty] | mm | inch
1x1 | 20 | 0 | 1177.226 | 1.444 | 0.066541 | 0.78634 | 0.0005 | 1337.002 | 1.523 | 0.070182 | 0.67481 | 0.0005 | 2.8462 | 0.13117 | 2.5762 | 0.11873 | 1 | 0.065175 | 1 | 0.065177 | 1 | 0.046086 | 1 | 1 | 0.99365 | 0.86546 | 0.70503 | 0.54665 | 0.44214 | x | y0 | y1 | y2 | y3 | y4 | y5 | x | y15 | y25 | y35 | y45 | y55 | y65 | r | 28.5 | 1.122 | 0 |  |  |  |  | 0
2x1 | 30 | 0 | 1177.245 | 1.436 | 0.066173 | 0.56116 | 0.0005 | 1337.006 | 1.503 | 0.06926 | 0.49674 | 0.0005 | 2.0199 | 0.093097 | 1.8714 | 0.086259 | 0.70968 | 0.046256 | 0.72645 | 0.04735 | 0.71807 | 0.033097 | 2 | 0.71807 | 0.70166 | 0.5995 | 0.54081 | 0.4664 | 0.38768 | [14 empty] | z | 20 | 0.7874 | 0.1 |  |  |  |  | 0
3x1 | 40 | 0 | 1177.213 | 1.39 | 0.064053 | 0.42818 | 0.0005 | 1336.998 | 1.45 | 0.066818 | 0.37672 | 0.0005 | 1.4919 | 0.068769 | 1.3692 | 0.063121 | 0.52416 | 0.034166 | 0.5315 | 0.034646 | 0.52783 | 0.024329 | 3 | 0.52783 | 0.50681 | 0.49422 | 0.41773 | 0.37284 | 0.31873 | 0 | 1.809 | 1.8276 | 1.4764 | 1.1216 | 0.77438 | 0.59561 | 0 | 1 | [5 empty] | h | 30 | 1.1811 | 0.2 |  |  |  |  | 0
4x1 | 50 | 0 | 1177.201 | 1.46 | 0.067279 | 0.3188 | 0.0005 | 1336.947 | 1.49 | 0.068661 | 0.27712 | 0.0005 | 1.1667 | 0.053794 | 1.035 | 0.047731 | 0.40991 | 0.026722 | 0.40177 | 0.026194 | 0.40584 | 0.01871 | 4 | 0.40584 | 0.39358 | 0.3725 | 0.30125 | 0.30962 | 0.27309 | 0.5 | 1.7864 | 1.8042 | 1.4596 | 1.1104 | 0.76843 | 0.59164 | 0.5 | 1.0008 | [5 empty] | h1 | 50 | 1.9685 | 0.3 |  |  |  |  | 0
5x1 | 60 | 0 | 1177.209 | 1.397 | 0.064375 | 0.25979 | 0.0005 | 1336.957 | 1.59 | 0.073269 | 0.2052 | 0.0005 | 0.90973 | 0.041958 | 0.81784 | 0.03774 | 0.31963 | 0.02084 | 0.31746 | 0.020705 | 0.31855 | 0.014688 | 5 | 0.31855 | 0.32355 | 0.28504 | 0.24362 | 0.25103 | 0.21406 | 1 | 1.7638 | 1.7808 | 1.4428 | 1.0991 | 0.76247 | 0.58767 | 1 | 1.0013 | [5 empty] | h2 | 20 | 0.7874 | 0.4 |  |  |  |  | 0
6x1 | 70 | 0 | 1177.332 | 1.41 | 0.064974 | 0.20827 | 0.0005 | 1337.112 | 1.31 | 0.060366 | 0.18874 | 0.0005 | 0.73612 | 0.033967 | 0.61975 | 0.028606 | 0.25863 | 0.016867 | 0.24057 | 0.015692 | 0.2496 | 0.011519 | 6 | 0.2496 | 0.25532 | 0.2414 | 0.22208 | 0.20578 | 0.1963 | 1.5 | 1.7412 | 1.7574 | 1.4261 | 1.0879 | 0.75652 | 0.5837 | 1.5 | 1.0017 | [8 empty] | 0.5 |  |  |  |  | 0
1x2 | 20 | 10 | 1177.409 | 1.434 | 0.06608 | 0.7879 | 0.0005 | 1337.155 | 1.491 | 0.068707 | 0.68395 | 0.0005 | 2.8321 | 0.13052 | 2.5562 | 0.11781 | 0.99505 | 0.064852 | 0.99225 | 0.064672 | 0.99365 | 0.045794 | [7 empty] | 2 | 1.7187 | 1.734 | 1.4094 | 1.0767 | 0.75057 | 0.57973 | 2 | 1.0018 | [5 empty] | r_corr |  | 1 | 0.6 |  |  |  |  | 0
2x2 | 30 | 10 | 1177.452 | 1.446 | 0.066633 | 0.5418 | 0.0005 | 1337.31 | 1.476 | 0.068016 | 0.4967 | 0.0005 | 1.9638 | 0.090513 | 1.8377 | 0.084703 | 0.68997 | 0.044971 | 0.71334 | 0.046496 | 0.70166 | 0.032343 | [7 empty] | 2.5 | 1.6962 | 1.7107 | 1.3926 | 1.0656 | 0.74462 | 0.57577 | 2.5 | 1.0017 | [5 empty] | z_corr |  | 0.70175 | 0.7 |  |  |  |  | 0
3x2 | 40 | 10 | 1177.522 | 1.41 | 0.064974 | 0.4033 | 0.0005 | 1337.335 | 1.5 | 0.069122 | 0.35136 | 0.0005 | 1.4254 | 0.065708 | 1.3211 | 0.060906 | 0.5008 | 0.032644 | 0.51281 | 0.033429 | 0.50681 | 0.023362 | [7 empty] | 3 | 1.6737 | 1.6875 | 1.376 | 1.0544 | 0.73868 | 0.5718 | 3 | 1.0013 | [5 empty] | h_corr |  | 1.0526 | 0.8 |  |  |  |  | 0
4x2 | 50 | 10 | 1178.299 | 1.403 | 0.064652 | 0.3184 | 0.0005 | 1338.269 | 1.497 | 0.068984 | 0.27031 | 0.0005 | 1.1197 | 0.051629 | 1.0143 | 0.046779 | 0.39342 | 0.025647 | 0.39374 | 0.025671 | 0.39358 | 0.018144 | [7 empty] | 3.5 | 1.6513 | 1.6643 | 1.3593 | 1.0433 | 0.73274 | 0.56784 | 3.5 | 1.0008 | [5 empty] | h1_corr |  | 1.7544 | 0.9 |  |  |  |  | 0
5x2 | 60 | 10 | 1178.362 | 1.37 | 0.063131 | 0.25983 | 0.0005 | 1338.322 | 1.47 | 0.067739 | 0.23324 | 0.0005 | 0.89227 | 0.041153 | 0.85943 | 0.039647 | 0.3135 | 0.02044 | 0.33361 | 0.021754 | 0.32355 | 0.014925 | [7 empty] | 4 | 1.629 | 1.6412 | 1.3427 | 1.0321 | 0.72682 | 0.56389 | 4 | 1 | [5 empty] | h2_corr |  | 0.70175 | 0.95 |  |  |  |  | 0
6x2 | 70 | 10 | 1177.9 | 1.49 | 0.068661 | 0.19746 | 0.0005 | 1337.807 | 1.49 | 0.068661 | 0.17349 | 0.0005 | 0.73749 | 0.034036 | 0.64796 | 0.029917 | 0.25911 | 0.0169 | 0.25152 | 0.016409 | 0.25532 | 0.011777 | [7 empty] | 4.5 | 1.6068 | 1.6181 | 1.3262 | 1.021 | 0.7209 | 0.55994 | 4.5 | 0.999 | [8 empty] | 1 |  |  |  |  | 0
1x3 | 20 | 20 | 1178.027 | 1.458 | 0.067186 | 0.67462 | 0.0005 | 1337.925 | 1.498 | 0.06903 | 0.59323 | 0.0005 | 2.4655 | 0.11363 | 2.2275 | 0.10266 | 0.86624 | 0.056458 | 0.86467 | 0.056358 | 0.86546 | 0.039887 | [7 empty] | 5 | 1.5846 | 1.5952 | 1.3097 | 1.01 | 0.71499 | 0.55599 | 5 | 0.99778 | [8 empty] | 1.1 |  |  |  |  | 0
2x3 | 30 | 20 | 1177.802 | 1.435 | 0.066126 | 0.47282 | 0.0005 | 1337.617 | 1.523 | 0.070182 | 0.40587 | 0.0005 | 1.7007 | 0.078392 | 1.5494 | 0.071425 | 0.59754 | 0.038948 | 0.60145 | 0.039205 | 0.5995 | 0.027631 | [7 empty] | 5.5 | 1.5626 | 1.5724 | 1.2933 | 0.99894 | 0.70909 | 0.55205 | 5.5 | 0.99635 | [8 empty] | 1.2 |  |  |  |  | 0
3x3 | 40 | 20 | 1177.895 | 1.398 | 0.064421 | 0.40067 | 0.0005 | 1337.668 | 1.49 | 0.068661 | 0.34153 | 0.0005 | 1.404 | 0.064724 | 1.2756 | 0.058809 | 0.4933 | 0.032155 | 0.49514 | 0.032278 | 0.49422 | 0.022781 | [7 empty] | 6 | 1.5407 | 1.5497 | 1.2769 | 0.98794 | 0.7032 | 0.54811 | 6 | 0.99469 | [8 empty] | 1.3 |  |  |  |  | 0
4x3 | 50 | 20 | 1177.813 | 1.49 | 0.068661 | 0.28832 | 0.0005 | 1337.678 | 1.52 | 0.070043 | 0.2479 | 0.0005 | 1.0769 | 0.049658 | 0.94453 | 0.043567 | 0.37835 | 0.024666 | 0.36665 | 0.023907 | 0.3725 | 0.017175 | [7 empty] | 6.5 | 1.5189 | 1.5272 | 1.2606 | 0.97699 | 0.69733 | 0.54418 | 6.5 | 0.99282 | [8 empty] | 1.4 |  |  |  |  | 0
5x3 | 60 | 20 | 1177.846 | 1.49 | 0.068661 | 0.22264 | 0.0005 | 1337.67 | 1.508 | 0.06949 | 0.18941 | 0.0005 | 0.83153 | 0.038363 | 0.71597 | 0.033047 | 0.29215 | 0.019051 | 0.27792 | 0.018128 | 0.28504 | 0.013149 | [7 empty] | 7 | 1.4972 | 1.5048 | 1.2444 | 0.96607 | 0.69147 | 0.54026 | 7 | 0.99073 | [8 empty] | 1.5 |  |  |  |  | 0
6x3 | 70 | 20 | 1177.848 | 1.49 | 0.068661 | 0.185 | 0.0005 | 1337.675 | 1.49 | 0.068661 | 0.16557 | 0.0005 | 0.69094 | 0.031894 | 0.6184 | 0.028558 | 0.24276 | 0.015835 | 0.24005 | 0.015661 | 0.2414 | 0.011136 | [7 empty] | 7.5 | 1.4756 | 1.4825 | 1.2283 | 0.9552 | 0.68562 | 0.53634 | 7.5 | 0.98842 | [8 empty] | 1.6 |  |  |  |  | 0
1x4 | 20 | 30 | 1177.785 | 1.44 | 0.066357 | 0.55054 | 0.0005 | 1337.64 | 1.493 | 0.068799 | 0.49004 | 0.0005 | 1.9872 | 0.09159 | 1.8339 | 0.084529 | 0.69819 | 0.045506 | 0.71188 | 0.046401 | 0.70503 | 0.032496 | [7 empty] | 8 | 1.4542 | 1.4604 | 1.2123 | 0.94438 | 0.67979 | 0.53243 | 8 | 0.9859 | [8 empty] | 1.7 |  |  |  |  | 0
2x4 | 30 | 30 | 1177.997 | 1.426 | 0.065712 | 0.42828 | 0.0005 | 1337.916 | 1.508 | 0.06949 | 0.37058 | 0.0005 | 1.5309 | 0.070567 | 1.4008 | 0.064578 | 0.53786 | 0.035059 | 0.54376 | 0.035446 | 0.54081 | 0.024928 | [7 empty] | 8.5 | 1.433 | 1.4384 | 1.1964 | 0.9336 | 0.67397 | 0.52853 | 8.5 | 0.98317 | [8 empty] | 1.8 |  |  |  |  | 0
3x4 | 40 | 30 | 1178.002 | 1.41 | 0.064974 | 0.32431 | 0.0005 | 1337.873 | 1.59 | 0.073269 | 0.27972 | 0.0005 | 1.1462 | 0.052849 | 1.1148 | 0.051411 | 0.40272 | 0.026253 | 0.43275 | 0.028214 | 0.41773 | 0.01927 | [7 empty] | 9 | 1.4119 | 1.4167 | 1.1805 | 0.92288 | 0.66817 | 0.52464 | 9 | 0.98023 | [8 empty] | 1.9 |  |  |  |  | 0
4x4 | 50 | 30 | 1178.1 | 1.41 | 0.064974 | 0.23948 | 0.0005 | 1337.872 | 1.59 | 0.073269 | 0.19722 | 0.0005 | 0.8464 | 0.039043 | 0.78602 | 0.036275 | 0.29738 | 0.019391 | 0.30511 | 0.0199 | 0.30125 | 0.013893 | [7 empty] | 9.5 | 1.391 | 1.3951 | 1.1648 | 0.91222 | 0.66239 | 0.52076 | 9.5 | 0.97709 | [8 empty] | 2 |  |  |  |  | 0
5x4 | 60 | 30 | 1178.331 | 1.4 | 0.064514 | 0.20206 | 0.0005 | 1338.169 | 1.31 | 0.060366 | 0.1868 | 0.0005 | 0.70907 | 0.032722 | 0.6134 | 0.028314 | 0.24913 | 0.016248 | 0.23811 | 0.015531 | 0.24362 | 0.011238 | [7 empty] | 10 | 1.3703 | 1.3737 | 1.1492 | 0.90162 | 0.65663 | 0.51689 | 10 | 0.97373 | [8 empty] | 2.1 |  |  |  |  | 0
6x4 | 70 | 30 | 1178.358 | 1.31 | 0.060366 | 0.1841 | 0.0005 | 1338.249 | 1.52 | 0.070043 | 0.1567 | 0.0005 | 0.60454 | 0.027906 | 0.59706 | 0.027579 | 0.2124 | 0.013855 | 0.23176 | 0.015123 | 0.22208 | 0.010255 | [7 empty] | 10.5 | 1.3497 | 1.3525 | 1.1337 | 0.89107 | 0.65089 | 0.51303 | 10.5 | 0.97018 | [8 empty] | 2.2 |  |  |  |  | 0
1x5 | 20 | 40 | 1178.412 | 1.467 | 0.067601 | 0.42832 | 0.0005 | 1338.275 | 1.465 | 0.067509 | 0.37877 | 0.0005 | 1.575 | 0.072602 | 1.3909 | 0.064121 | 0.55338 | 0.03607 | 0.53992 | 0.035195 | 0.54665 | 0.025198 | [7 empty] | 11 | 1.3293 | 1.3315 | 1.1183 | 0.88059 | 0.64517 | 0.50918 | 11 | 0.96642 | [8 empty] | 2.3 |  |  |  |  | 0
2x5 | 30 | 40 | 1178.412 | 1.435 | 0.066126 | 0.36785 | 0.0005 | 1338.221 | 1.488 | 0.068569 | 0.32318 | 0.0005 | 1.3232 | 0.060999 | 1.2054 | 0.055579 | 0.46488 | 0.030304 | 0.46792 | 0.030504 | 0.4664 | 0.021499 | [7 empty] | 11.5 | 1.3092 | 1.3107 | 1.1031 | 0.87018 | 0.63947 | 0.50534 | 11.5 | 0.96246 | [8 empty] | 2.4 |  |  |  |  | 0
3x5 | 40 | 40 | 1178.477 | 1.496 | 0.068937 | 0.28127 | 0.0005 | 1338.376 | 1.537 | 0.070827 | 0.25082 | 0.0005 | 1.0548 | 0.04864 | 0.96633 | 0.044571 | 0.37058 | 0.02416 | 0.3751 | 0.024458 | 0.37284 | 0.01719 | [7 empty] | 12 | 1.2892 | 1.2901 | 1.0879 | 0.85983 | 0.63379 | 0.50151 | 12 | 0.9583 | [8 empty] | 2.5 |  |  |  |  | 0
4x5 | 50 | 40 | 1178.412 | 1.39 | 0.064053 | 0.24539 | 0.0005 | 1338.309 | 1.482 | 0.068292 | 0.22112 | 0.0005 | 0.85497 | 0.039437 | 0.82143 | 0.037898 | 0.30039 | 0.019587 | 0.31886 | 0.020793 | 0.30962 | 0.014283 | [7 empty] | 12.5 | 1.2694 | 1.2698 | 1.0729 | 0.84955 | 0.62814 | 0.49769 | 12.5 | 0.95396 | [8 empty] | 2.6 |  |  |  |  | 0
5x5 | 60 | 40 | 1178.361 | 1.41 | 0.064974 | 0.21148 | 0.0005 | 1338.383 | 1.47 | 0.067739 | 0.1674 | 0.0005 | 0.74745 | 0.034489 | 0.61683 | 0.028484 | 0.26261 | 0.017126 | 0.23944 | 0.015621 | 0.25103 | 0.01159 | [7 empty] | 13 | 1.2499 | 1.2497 | 1.0581 | 0.83935 | 0.62251 | 0.49389 | 13 | 0.94942 | [8 empty] | 2.7 |  |  |  |  | 0
6x5 | 70 | 40 | 1178.409 | 1.31 | 0.060366 | 0.1734 | 0.0005 | 1338.27 | 1.31 | 0.060366 | 0.16594 | 0.0005 | 0.56938 | 0.026289 | 0.54488 | 0.025162 | 0.20005 | 0.01305 | 0.21151 | 0.013799 | 0.20578 | 0.0094965 | [7 empty] | 13.5 | 1.2305 | 1.2298 | 1.0433 | 0.82922 | 0.61691 | 0.49009 | 13.5 | 0.94469 | [8 empty] | 2.8 |  |  |  |  | 0
1x6 | 20 | 50 | 1178.784 | 1.49 | 0.068661 | 0.33378 | 0.0005 | 1338.716 | 1.49 | 0.068661 | 0.30783 | 0.0005 | 1.2466 | 0.057476 | 1.1497 | 0.053013 | 0.43799 | 0.028552 | 0.44629 | 0.029095 | 0.44214 | 0.020382 | [7 empty] | 14 | 1.2114 | 1.2101 | 1.0287 | 0.81916 | 0.61133 | 0.48632 | 14 | 0.93978 | [8 empty] | 2.9 |  |  |  |  | 0
2x6 | 30 | 50 | 1178.896 | 1.49 | 0.068661 | 0.28937 | 0.0005 | 1338.852 | 1.59 | 0.073269 | 0.25574 | 0.0005 | 1.0808 | 0.049838 | 1.0193 | 0.047011 | 0.37972 | 0.024756 | 0.39565 | 0.025797 | 0.38768 | 0.017877 | [7 empty] | 14.5 | 1.1925 | 1.1907 | 1.0143 | 0.80919 | 0.60578 | 0.48255 | 14.5 | 0.93469 | [8 empty] | 3 |  |  |  |  | 0
3x6 | 40 | 50 | 1178.928 | 1.49 | 0.068661 | 0.23954 | 0.0005 | 1339.007 | 1.69 | 0.077877 | 0.19651 | 0.0005 | 0.89466 | 0.041269 | 0.83245 | 0.038419 | 0.31433 | 0.020496 | 0.32314 | 0.021076 | 0.31873 | 0.014699 | [7 empty] | 15 | 1.1738 | 1.1715 | 1 | 0.79929 | 0.60025 | 0.4788 | 15 | 0.92942 | [8 empty] | 3.1 |  |  |  |  | 0
4x6 | 50 | 50 | 1178.978 | 1.49 | 0.068661 | 0.2113 | 0.0005 | 1338.939 | 1.59 | 0.073269 | 0.17381 | 0.0005 | 0.78919 | 0.036415 | 0.69272 | 0.031983 | 0.27728 | 0.018082 | 0.2689 | 0.017542 | 0.27309 | 0.012597 | [7 empty] | 15.5 | 1.1554 | 1.1526 | 0.98585 | 0.78948 | 0.59476 | 0.47506 | 15.5 | 0.92397 | [8 empty] | 3.2 |  |  |  |  | 0
5x6 | 60 | 50 | 1179.051 | 1.21 | 0.055758 | 0.19409 | 0.0005 | 1338.991 | 1.59 | 0.073269 | 0.14304 | 0.0005 | 0.58868 | 0.027169 | 0.5701 | 0.026346 | 0.20683 | 0.01349 | 0.2213 | 0.014443 | 0.21406 | 0.0098817 | [7 empty] | 16 | 1.1372 | 1.1339 | 0.97185 | 0.77975 | 0.58929 | 0.47134 | 16 | 0.91835 | [8 empty] | 3.3 |  |  |  |  | 0
6x6 | 70 | 50 | 1179.045 | 1.59 | 0.073269 | 0.14017 | 0.0005 | 1339.098 | 1.69 | 0.077877 | 0.11938 | 0.0005 | 0.55866 | 0.025821 | 0.50573 | 0.023401 | 0.19628 | 0.012811 | 0.19631 | 0.012821 | 0.1963 | 0.0090621 | [7 empty] | 16.5 | 1.1192 | 1.1155 | 0.958 | 0.7701 | 0.58385 | 0.46763 | 16.5 | 0.91256 | [8 empty] | 3.4 |  |  |  |  | 0
[30 empty] | 17 | 1.1014 | 1.0973 | 0.9443 | 0.76054 | 0.57844 | 0.46394 | 17 | 0.90661 | [8 empty] | 3.5 |  |  |  |  | 0
|  |  |  | 0.065941 | [25 empty] | 17.5 | 1.0839 | 1.0794 | 0.93076 | 0.75106 | 0.57306 | 0.46027 | 17.5 | 0.9005 | [8 empty] | 3.6 |  |  |  |  | 0
|  |  |  | 1.431 | [25 empty] | 18 | 1.0667 | 1.0617 | 0.91738 | 0.74168 | 0.56772 | 0.45661 | 18 | 0.89423 | [8 empty] | 3.7 |  |  |  |  | 0
[30 empty] | 18.5 | 1.0496 | 1.0443 | 0.90416 | 0.73238 | 0.5624 | 0.45297 | 18.5 | 0.88781 | [8 empty] | 3.8 |  |  |  |  | 0
|  |  |  | 0.046081 | [25 empty] | 19 | 1.0328 | 1.0272 | 0.8911 | 0.72317 | 0.55712 | 0.44934 | 19 | 0.88124 | [8 empty] | 3.9 |  |  |  |  | 0
[30 empty] | 19.5 | 1.0163 | 1.0103 | 0.8782 | 0.71406 | 0.55186 | 0.44573 | 19.5 | 0.87452 | [8 empty] | 4 |  |  |  |  | 0
[30 empty] | 20 | 1 | 0.99365 | 0.86546 | 0.70503 | 0.54665 | 0.44214 | 20 | 0.86766
[30 empty] | 20.5 | 0.98394 | 0.97727 | 0.85288 | 0.6961 | 0.54146 | 0.43857 | 20.5 | 0.86067
[30 empty] | 21 | 0.96812 | 0.96115 | 0.84047 | 0.68726 | 0.53631 | 0.43501 | 21 | 0.85354
[30 empty] | 21.5 | 0.95255 | 0.94528 | 0.82822 | 0.67852 | 0.5312 | 0.43148 | 21.5 | 0.84629
[30 empty] | 22 | 0.93722 | 0.92968 | 0.81613 | 0.66987 | 0.52612 | 0.42796 | 22 | 0.8389
[30 empty] | 22.5 | 0.92213 | 0.91432 | 0.80421 | 0.66132 | 0.52107 | 0.42446 | 22.5 | 0.8314
[30 empty] | 23 | 0.90727 | 0.89922 | 0.79245 | 0.65286 | 0.51606 | 0.42098 | 23 | 0.82379
[30 empty] | 23.5 | 0.89266 | 0.88437 | 0.78085 | 0.6445 | 0.51109 | 0.41752 | 23.5 | 0.81606
[30 empty] | 24 | 0.87829 | 0.86977 | 0.76942 | 0.63623 | 0.50616 | 0.41408 | 24 | 0.80823
[30 empty] | 24.5 | 0.86415 | 0.85542 | 0.75815 | 0.62806 | 0.50126 | 0.41066 | 24.5 | 0.80029
[30 empty] | 25 | 0.85024 | 0.84132 | 0.74704 | 0.61999 | 0.49639 | 0.40726 | 25 | 0.79226
[30 empty] | 25.5 | 0.83657 | 0.82746 | 0.73609 | 0.61201 | 0.49157 | 0.40388 | 25.5 | 0.78413
[30 empty] | 26 | 0.82312 | 0.81384 | 0.72531 | 0.60413 | 0.48678 | 0.40052 | 26 | 0.77592
[30 empty] | 26.5 | 0.80991 | 0.80046 | 0.71468 | 0.59635 | 0.48203 | 0.39718 | 26.5 | 0.76762
[30 empty] | 27 | 0.79692 | 0.78732 | 0.70422 | 0.58866 | 0.47732 | 0.39386 | 27 | 0.75924
[30 empty] | 27.5 | 0.78416 | 0.77442 | 0.69391 | 0.58107 | 0.47265 | 0.39057 | 27.5 | 0.75079
[30 empty] | 28 | 0.77161 | 0.76174 | 0.68375 | 0.57357 | 0.46802 | 0.38729 | 28 | 0.74227
[30 empty] | 28.5 | 0.75929 | 0.7493 | 0.67376 | 0.56617 | 0.46342 | 0.38404 | 28.5 | 0.73368
[30 empty] | 29 | 0.74718 | 0.73708 | 0.66391 | 0.55886 | 0.45887 | 0.3808 | 29 | 0.72503
[30 empty] | 29.5 | 0.73529 | 0.72508 | 0.65422 | 0.55165 | 0.45435 | 0.37759 | 29.5 | 0.71632
[30 empty] | 30 | 0.72361 | 0.7133 | 0.64469 | 0.54453 | 0.44987 | 0.3744 | 30 | 0.70756
[30 empty] | 30.5 | 0.71213 | 0.70174 | 0.6353 | 0.5375 | 0.44543 | 0.37124 | 30.5 | 0.69876
[30 empty] | 31 | 0.70086 | 0.6904 | 0.62606 | 0.53057 | 0.44103 | 0.36809 | 31 | 0.68991
[30 empty] | 31.5 | 0.6898 | 0.67926 | 0.61697 | 0.52373 | 0.43667 | 0.36497 | 31.5 | 0.68102
[30 empty] | 32 | 0.67893 | 0.66833 | 0.60802 | 0.51699 | 0.43235 | 0.36187 | 32 | 0.6721
[30 empty] | 32.5 | 0.66826 | 0.65761 | 0.59922 | 0.51033 | 0.42807 | 0.35879 | 32.5 | 0.66314
[30 empty] | 33 | 0.65779 | 0.64709 | 0.59056 | 0.50376 | 0.42383 | 0.35573 | 33 | 0.65417
[30 empty] | 33.5 | 0.64751 | 0.63676 | 0.58204 | 0.49729 | 0.41962 | 0.3527 | 33.5 | 0.64517
[30 empty] | 34 | 0.63741 | 0.62663 | 0.57366 | 0.4909 | 0.41546 | 0.34969 | 34 | 0.63615
[30 empty] | 34.5 | 0.6275 | 0.61669 | 0.56541 | 0.4846 | 0.41133 | 0.3467 | 34.5 | 0.62713
[30 empty] | 35 | 0.61777 | 0.60694 | 0.5573 | 0.47839 | 0.40725 | 0.34373 | 35 | 0.61809
[30 empty] | 35.5 | 0.60822 | 0.59737 | 0.54933 | 0.47227 | 0.4032 | 0.34079 | 35.5 | 0.60905
[30 empty] | 36 | 0.59885 | 0.58798 | 0.54149 | 0.46623 | 0.39919 | 0.33787 | 36 | 0.60001
[30 empty] | 36.5 | 0.58965 | 0.57878 | 0.53377 | 0.46028 | 0.39523 | 0.33497 | 36.5 | 0.59097
[30 empty] | 37 | 0.58062 | 0.56974 | 0.52619 | 0.45441 | 0.3913 | 0.3321 | 37 | 0.58194
[30 empty] | 37.5 | 0.57176 | 0.56088 | 0.51873 | 0.44863 | 0.38741 | 0.32924 | 37.5 | 0.57292
[30 empty] | 38 | 0.56306 | 0.55219 | 0.51139 | 0.44293 | 0.38355 | 0.32641 | 38 | 0.56392
[30 empty] | 38.5 | 0.55452 | 0.54366 | 0.50418 | 0.43731 | 0.37974 | 0.32361 | 38.5 | 0.55493
[30 empty] | 39 | 0.54614 | 0.5353 | 0.49709 | 0.43177 | 0.37597 | 0.32082 | 39 | 0.54597
[30 empty] | 39.5 | 0.53792 | 0.52709 | 0.49012 | 0.42631 | 0.37223 | 0.31806 | 39.5 | 0.53704
[30 empty] | 40 | 0.52985 | 0.51905 | 0.48326 | 0.42093 | 0.36853 | 0.31532 | 40 | 0.52813
[30 empty] | 40.5 | 0.52193 | 0.51115 | 0.47652 | 0.41563 | 0.36487 | 0.31261 | 40.5 | 0.51926
[30 empty] | 41 | 0.51416 | 0.50341 | 0.4699 | 0.41041 | 0.36125 | 0.30991 | 41 | 0.51042
[30 empty] | 41.5 | 0.50653 | 0.49581 | 0.46338 | 0.40526 | 0.35766 | 0.30724 | 41.5 | 0.50163
[30 empty] | 42 | 0.49904 | 0.48835 | 0.45698 | 0.40019 | 0.35411 | 0.30459 | 42 | 0.49287
[30 empty] | 42.5 | 0.49169 | 0.48104 | 0.45068 | 0.39519 | 0.3506 | 0.30197 | 42.5 | 0.48417
[30 empty] | 43 | 0.48448 | 0.47387 | 0.44449 | 0.39026 | 0.34713 | 0.29936 | 43 | 0.47551
[30 empty] | 43.5 | 0.4774 | 0.46683 | 0.43841 | 0.38541 | 0.34369 | 0.29678 | 43.5 | 0.4669
[30 empty] | 44 | 0.47045 | 0.45993 | 0.43243 | 0.38063 | 0.34029 | 0.29422 | 44 | 0.45835
[30 empty] | 44.5 | 0.46363 | 0.45315 | 0.42654 | 0.37592 | 0.33692 | 0.29168 | 44.5 | 0.44986
[30 empty] | 45 | 0.45694 | 0.44651 | 0.42076 | 0.37128 | 0.3336 | 0.28917 | 45 | 0.44143
[30 empty] | 45.5 | 0.45037 | 0.43999 | 0.41508 | 0.36671 | 0.3303 | 0.28668 | 45.5 | 0.43306
[30 empty] | 46 | 0.44392 | 0.43359 | 0.40949 | 0.36221 | 0.32705 | 0.28421 | 46 | 0.42475
[30 empty] | 46.5 | 0.43759 | 0.42731 | 0.404 | 0.35777 | 0.32382 | 0.28176 | 46.5 | 0.41651
[30 empty] | 47 | 0.43138 | 0.42116 | 0.39859 | 0.3534 | 0.32064 | 0.27933 | 47 | 0.40835
[30 empty] | 47.5 | 0.42528 | 0.41511 | 0.39328 | 0.34909 | 0.31748 | 0.27693 | 47.5 | 0.40025
[30 empty] | 48 | 0.41929 | 0.40918 | 0.38806 | 0.34485 | 0.31437 | 0.27454 | 48 | 0.39223
[30 empty] | 48.5 | 0.41341 | 0.40336 | 0.38293 | 0.34067 | 0.31128 | 0.27218 | 48.5 | 0.38429
[30 empty] | 49 | 0.40764 | 0.39765 | 0.37788 | 0.33655 | 0.30823 | 0.26984 | 49 | 0.37642
[30 empty] | 49.5 | 0.40198 | 0.39204 | 0.37292 | 0.3325 | 0.30522 | 0.26752 | 49.5 | 0.36864
[30 empty] | 50 | 0.39641 | 0.38654 | 0.36804 | 0.3285 | 0.30223 | 0.26522 | 50 | 0.36093
[30 empty] | 50.5 | 0.39095 | 0.38114 | 0.36325 | 0.32457 | 0.29928 | 0.26295 | 50.5 | 0.35331
[30 empty] | 51 | 0.38559 | 0.37584 | 0.35853 | 0.32069 | 0.29637 | 0.26069 | 51 | 0.34578
[30 empty] | 51.5 | 0.38032 | 0.37063 | 0.35389 | 0.31687 | 0.29348 | 0.25846 | 51.5 | 0.33833
[30 empty] | 52 | 0.37515 | 0.36553 | 0.34933 | 0.31311 | 0.29063 | 0.25624 | 52 | 0.33097
[30 empty] | 52.5 | 0.37007 | 0.36051 | 0.34485 | 0.3094 | 0.28781 | 0.25405 | 52.5 | 0.3237
[30 empty] | 53 | 0.36508 | 0.35559 | 0.34044 | 0.30575 | 0.28502 | 0.25188 | 53 | 0.31652
[30 empty] | 53.5 | 0.36018 | 0.35075 | 0.33611 | 0.30215 | 0.28226 | 0.24972 | 53.5 | 0.30943
[30 empty] | 54 | 0.35537 | 0.34601 | 0.33184 | 0.29861 | 0.27954 | 0.24759 | 54 | 0.30244
[30 empty] | 54.5 | 0.35064 | 0.34135 | 0.32765 | 0.29512 | 0.27684 | 0.24548 | 54.5 | 0.29553
[30 empty] | 55 | 0.346 | 0.33677 | 0.32352 | 0.29168 | 0.27418 | 0.24339 | 55 | 0.28873
[30 empty] | 55.5 | 0.34144 | 0.33228 | 0.31947 | 0.28829 | 0.27154 | 0.24131 | 55.5 | 0.28202
[30 empty] | 56 | 0.33696 | 0.32786 | 0.31548 | 0.28495 | 0.26894 | 0.23926 | 56 | 0.2754
[30 empty] | 56.5 | 0.33256 | 0.32353 | 0.31155 | 0.28166 | 0.26636 | 0.23723 | 56.5 | 0.26888
[30 empty] | 57 | 0.32823 | 0.31927 | 0.30769 | 0.27842 | 0.26381 | 0.23522 | 57 | 0.26246
[30 empty] | 57.5 | 0.32399 | 0.31509 | 0.3039 | 0.27523 | 0.2613 | 0.23322 | 57.5 | 0.25614
[30 empty] | 58 | 0.31981 | 0.31098 | 0.30016 | 0.27208 | 0.25881 | 0.23125 | 58 | 0.24991
[30 empty] | 58.5 | 0.31571 | 0.30695 | 0.29649 | 0.26898 | 0.25635 | 0.22929 | 58.5 | 0.24379
[30 empty] | 59 | 0.31168 | 0.30299 | 0.29288 | 0.26593 | 0.25392 | 0.22735 | 59 | 0.23776
[30 empty] | 59.5 | 0.30772 | 0.29909 | 0.28932 | 0.26292 | 0.25151 | 0.22543 | 59.5 | 0.23183
[30 empty] | 60 | 0.30382 | 0.29526 | 0.28582 | 0.25996 | 0.24913 | 0.22353 | 60 | 0.226
[30 empty] | 60.5 | 0.3 | 0.29151 | 0.28238 | 0.25704 | 0.24678 | 0.22165 | 60.5 | 0.22027
[30 empty] | 61 | 0.29624 | 0.28781 | 0.27899 | 0.25416 | 0.24446 | 0.21979 | 61 | 0.21464
[30 empty] | 61.5 | 0.29254 | 0.28418 | 0.27566 | 0.25132 | 0.24217 | 0.21795 | 61.5 | 0.20911
[30 empty] | 62 | 0.28891 | 0.28061 | 0.27238 | 0.24853 | 0.2399 | 0.21612 | 62 | 0.20367
[30 empty] | 62.5 | 0.28533 | 0.27711 | 0.26916 | 0.24577 | 0.23765 | 0.21431 | 62.5 | 0.19834
[30 empty] | 63 | 0.28182 | 0.27366 | 0.26598 | 0.24306 | 0.23543 | 0.21252 | 63 | 0.1931
[30 empty] | 63.5 | 0.27837 | 0.27027 | 0.26286 | 0.24038 | 0.23324 | 0.21074 | 63.5 | 0.18796
[30 empty] | 64 | 0.27497 | 0.26694 | 0.25978 | 0.23775 | 0.23107 | 0.20899 | 64 | 0.18292
[30 empty] | 64.5 | 0.27164 | 0.26367 | 0.25676 | 0.23515 | 0.22893 | 0.20725 | 64.5 | 0.17798
[30 empty] | 65 | 0.26835 | 0.26045 | 0.25378 | 0.23259 | 0.22681 | 0.20553 | 65 | 0.17313
[30 empty] | 65.5 | 0.26512 | 0.25729 | 0.25085 | 0.23006 | 0.22472 | 0.20382 | 65.5 | 0.16838
[30 empty] | 66 | 0.26195 | 0.25417 | 0.24796 | 0.22758 | 0.22265 | 0.20213 | 66 | 0.16373
[30 empty] | 66.5 | 0.25883 | 0.25112 | 0.24512 | 0.22512 | 0.2206 | 0.20046 | 66.5 | 0.15917
[30 empty] | 67 | 0.25576 | 0.24811 | 0.24232 | 0.22271 | 0.21858 | 0.19881 | 67 | 0.15471
[30 empty] | 67.5 | 0.25273 | 0.24515 | 0.23957 | 0.22033 | 0.21658 | 0.19717 | 67.5 | 0.15033
[30 empty] | 68 | 0.24976 | 0.24224 | 0.23686 | 0.21798 | 0.2146 | 0.19554 | 68 | 0.14606
[30 empty] | 68.5 | 0.24684 | 0.23938 | 0.23419 | 0.21566 | 0.21265 | 0.19394 | 68.5 | 0.14187
[30 empty] | 69 | 0.24396 | 0.23656 | 0.23156 | 0.21338 | 0.21072 | 0.19235 | 69 | 0.13777
[30 empty] | 69.5 | 0.24113 | 0.23379 | 0.22897 | 0.21113 | 0.20881 | 0.19077 | 69.5 | 0.13377
[30 empty] | 70 | 0.23835 | 0.23107 | 0.22642 | 0.20891 | 0.20692 | 0.18921 | 70 | 0.12985

## Results
|  |  | HP whisky attenuation |  | HP bottle attenuation |  |  | Polyethylene bottle attenuation | [28 empty] | cm |  | cm |  | cm |  | Attenuated Background | [7 empty] | Background Subtracted | [7 empty] | GeoEff and Attenuation |  |  |  | Final Values of Bottles (Bq) |  |  |  | GeoEff and Attenuation |  |  |  | Uncorrected Values of Whisky (Bq) |  |  |  |  | GeoEff and Attenuation |  |  | Final Values of Whisky(Bq) | [6 empty] | WHISKY BOTTLE | [7 empty] | Total Activity of Each peak corrected for branching ratio | [8 empty] | WHISKY | [7 empty] | Total Activity of Each peak corrected for branching ratio
corrected energy |  |  | cm^2 g^-1 | cm^-1 | cm^2 g^-1 | cm^-1 |  | cm^2 g^-1 | cm^-1 | Peak Max | Isotope |  | HighlandPark | unc | HighlandParkBottle | unc | HighlandParkWhisky | unc | PolyethyleneBottle | unc | Background | unc | Energy Efficiency | max | min | unc | HighlandPark | unc | HighlandParkBottle | unc | HighlandParkWhisky | unc | PolyethyleneBottle | unc | Background |  | HP Bottle thickness | unc | Whisky thickness | unc | LDPE tickness | unc | HighlandPark | unc | HighlandParkBottle | unc | HighlandParkWhisky | unc | PolyethyleneBottle | unc | HighlandPark | unc | HighlandParkBottle | unc | HighlandParkWhisky | unc | PolyethyleneBottle | unc | HighlandParkBottle | PolyethyleneBottle | unc |  | HighlandParkBottle | PolyethyleneBottle | unc | unc | HighlandParkBottle | PolyethyleneBottle | unc |  | HighlandParkBottle | unc |  | PolyethyleneBottle | unc | HighlandParkBottle | PolyethyleneBottle | unc | HighlandParkBottle | PolyethyleneBottle | unc | unc |  |  |  |  | Branching Ratio | unc |  |  |  |  | unc | in bottle/Bq | [9 empty] | Branching Ratio | unc |  |  |  |  | unc | in bottle/Bq | [7 empty] | Glass Bottle |  | Whisky |  |  |  | Total Activity/Bq
[117 empty] | Peak Energy/keV | Isotopes Present | Peak Activity/Bq | Uncertainty | Peak Activity/Bq | Uncertainty | Decay Chain | Isotope | Glass Bottle | Whisky
46.365 | 0.046365 |  | 0.2367 | 0.26034 | 0.4133 | 0.16401 |  | 0.214 | 0.23135 | 47 | Pb210 |  | 0.0093252 | 0.00024277 | 0.0095519 | 0.00020081 | 0.0066625 | 0.00028383 | 0.0080706 | 0.00032663 | 0.0074554 | 0.00017418 | 0.10323 | 0.10931 | 0.093834 | 0.0077387 | 0.090334 | 0.0071686 | 0.09253 | 0.0072042 | 0.06454 | 0.0055649 | 0.078181 | 0.0066604 | 0.072221 | 0.0056709 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.062327 | 0.91404 | 0.0704 | 0.9131 | 0.062485 | 0.91436 | 0.071157 | 0.91436 | 0.028007 | 0.91407 | 0.02213 | 0.91313 | 0.0020547 | 0.91437 | 0.0070234 | 0.91438 | 5.9154 | 6.5822 | 0.01 |  | 0.13091 | 0.046229 | 5.4015 | 6.0187 | 8.189 | 9.6318 | 0.01 | 0.012021 | 0.012021 | 0.035993 | -0.002745 | 0 | 0.0070919 | 13.2434 | 15.7897 | 0.01 | 0.1592 | 0 | 0.47667 |  | 47 | Pb210 |  |  | 4.252 | 0.04 |  |  |  |  | 127.0343 | 3.0787 |  | Th-234 | 121.4506 |  |  | 47 | Pb210 |  |  | 4.252 | 0.04 |  |  |  |  | 11.2105 | 3.7442 |  | Th-234 | 19.8496 |  |  | 46.365 | Pb210 | 3.0787 | 127.0343 | 3.7442 | 11.2105 | U238 | Th-234 | 121.4506 | 19.8496
63.3511 | 0.063351 |  | 0.2028 | 0.22306 | 0.2584 | 0.10254 |  | 0.194 | 0.20973 | 64 | Th234 | Th232 | 0.0113 | 0.00026724 | 0.0085577 | 0.00019008 | 0.0076298 | 0.00030374 | 0.0075154 | 0.0003152 | 0.00866 | 0.00018773 | 0.089587 | 0.09077 | 0.086832 | 0.001969 | 0.12614 | 0.0040723 | 0.095523 | 0.0029848 | 0.085166 | 0.0038728 | 0.083889 | 0.0039722 | 0.096666 | 0.0029841 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.084369 | 0.91496 | 0.095096 | 0.91319 | 0.08444 | 0.91519 | 0.095369 | 0.91519 | 0.041768 | 0.91497 | 0.00042791 | 0.9132 | 0.00072617 | 0.9152 | -0.01148 | 0.9152 | 5.7876 | 6.5621 | 0.01 |  | 0.0024766 | 0 | 5.2852 |  | 7.824 | 9.3051 | 0.01 | 0.041451 | 0.041451 | 0.04177 | 0.00072617 | 0.00072617 | 0.00072617 | 11.5597 | 13.8451 | 0.01 | 0.47916 | 0.010054 | 0.48285 | 0.010054 | 64 | Th234 | Th232 |  | 3.75 | 0.08 | 0.259 | 0.015 |  |  |  | 1297.1818 |  | Ra-226 | 36.98 |  |  | 64 | Th234 | Th232 |  | 3.75 | 0.08 | 0.259 | 0.015 |  |  |  | 139.801 |  | Ra-226 | 6.0439 |  |  | 63.3511 | Th234, Th232 | 1297.1818 |  | 139.801 |  |  | Ra-226 | 36.98 | 6.0439
76.3405 | 0.076341 |  | 0.1895 | 0.20843 | 0.2124 | 0.084286 |  | 0.1845 | 0.19946 | 77 | Pb214 | Pb212 | 0.034507 | 0.000467 | 0.042904 | 0.0004256 | [6 empty] | 0.078501 | 0.078443 | 0.077563 | 0.00049785 | 0.43958 | 0.0065697 | 0.54655 | 0.0064349 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.43958 | 0.91289 | 0.54655 | 0.91289 | 0 | 0.91287 | 0 | 0.91287 | 5.75 | 6.5525 | 0.01 |  | 3.1426 | 0 | 5.2492 |  | 7.6895 | 9.1665 | 0.01 | 0.030884 | 0.030884 | 0.72614 | 0 | 0 | 0 | 10.9665 | 13.1134 | 0.01 | 0.33869 | 0 | 7.9633 |  | 77 | Pb214 | Pb212 |  | 10.47 | 0.2 | 16.9 | 0.3 |  |  |  | 25.8864 |  | Pb-214 | 20.8134 |  |  | 77 | Pb214 | Pb212 |  | 10.47 | 0.2 | 16.9 | 0.3 |  |  |  | 3.6474 |  | Pb-214 | 3.0585 |  |  | 76.3405 | Pb214, Pb212 | 25.8864 |  | 3.6474 |  |  | Pb-214 | 20.8134 | 3.0585
92.3275 | 0.092327 |  | 0.1784 | 0.19622 | 0.1827 | 0.0725 |  | 0.1756 | 0.18984 | 93 | Th234 |  | 0.032798 | 0.00045528 | 0.038104 | 0.00040109 | 0.019117 | 0.00048078 | 0.018342 | 0.00049242 | 0.018651 | 0.0002755 | 0.067302 | 0.066628 | 0.067356 | 0.00036372 | 0.48732 | 0.0072593 | 0.56617 | 0.0066991 | 0.28404 | 0.0073067 | 0.27254 | 0.0074633 | 0.27712 | 0.0043588 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.24417 | 0.92009 | 0.27389 | 0.91381 | 0.24412 | 0.92042 | 0.27374 | 0.92042 | 0.24315 | 0.92012 | 0.29228 | 0.91384 | 0.039923 | 0.92045 | -0.0012057 | 0.92045 | 5.7258 | 6.5435 | 0.01 |  | 1.6735 | 0 | 5.2325 |  | 7.5828 | 9.0457 | 0.01 | 0.022453 | 0.022453 | 0.39565 | 0.039923 | 0.039923 | 0.039923 | 10.5018 | 12.5161 | 0.01 | 0.2358 | 0.49968 | 4.155 | 0.49968 | 93 | Th234 |  |  | 2.18 | 0.19 | 2.15 | 0.19 |  |  |  | 121.4506 |  | Bi-214 | 24.064 |  |  | 93 | Th234 |  |  | 2.18 | 0.19 | 2.15 | 0.19 |  |  |  | 19.8496 |  | Bi-214 | 0.4513 |  |  | 92.3275 | Th234 | 121.4506 |  | 19.8496 |  |  | Bi-214 | 24.064 | 0.4513
186.251 | 0.18625 |  | 0.1436 | 0.15794 | 0.1293 | 0.05131 |  | 0.1435 | 0.15514 | 187 | Ra226 |  | 0.012953 | 0.00028611 | 0.014529 | 0.00024767 | 0.005538 | 0.00025877 | 0.0060017 | 0.00028167 | 0.0064197 | 0.00016163 | 0.035285 | 0.03441 | 0.036057 | 0.00082356 | 0.36709 | 0.011797 | 0.41177 | 0.011901 | 0.15695 | 0.0081977 | 0.17009 | 0.0089155 | 0.18194 | 0.0062463 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.16278 | 0.91897 | 0.18042 | 0.91351 | 0.16266 | 0.91904 | 0.18011 | 0.91904 | 0.20431 | 0.91904 | 0.23135 | 0.91358 | -0.0057135 | 0.91907 | -0.010021 | 0.91908 | 5.6824 | 6.5113 | 0.01 |  | 1.3146 | 0 | 5.1914 |  | 7.2671 | 8.6364 | 0.01 | 0.023407 | 0.023407 | 0.3274 | -0.0057135 | 0 | 0.0057135 | 9.1796 | 10.7309 | 0.01 | 0.21486 | 0 | 3.0054 |  | 187 | Ra226 |  |  | 3.555 | 0.019 |  |  |  |  | 146.0305 | 36.98 |  | Pb-210 | 3.0787 |  |  | 187 | Ra226 |  |  | 3.555 | 0.019 |  |  |  |  | 84.5408 | 6.0439 |  | Pb-210 | 3.7442 |  |  | 186.251 | Ra226 | 36.98 | 146.0305 | 6.0439 | 84.5408 |  | Pb-210 | 3.0787 | 3.7442
239.2078 | 0.23921 |  | 0.1319 | 0.14507 | 0.1173 | 0.046548 |  | 0.132 | 0.1427 | 240 | Pb212 | Ra224 | 0.034586 | 0.00046753 | 0.038737 | 0.0004044 | 0.0038875 | 0.00021681 |  |  | 0.0033818 | 0.00011731 | 0.027666 | 0.026926 | 0.028352 | 0.00071297 | 1.2501 | 0.03638 | 1.4002 | 0.038932 | 0.14051 | 0.0086329 | 0 |  | 0.12224 | 0.0052824 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.11001 | 0.91732 | 0.12131 | 0.9133 | 0.10992 | 0.91735 | 0.12111 | 0.91735 | 1.1401 | 0.91804 | 1.2789 | 0.91413 | 0.030592 | 0.91739 | -0.12111 | 0.91735 | 5.6727 | 6.4998 | 0.01 |  | 7.2547 | 0 | 5.1856 |  | 7.1593 | 8.4884 | 0.01 | 0.1268 | 0.1268 | 1.8313 | 0.030592 | 0.030592 | 0.030592 | 8.7621 | 10.163 | 0.01 | 1.111 | 0.31091 | 16.0459 | 0.31091 | 240 | Pb212 | Ra224 | Pb214 | 43.6 | 0.5 | 4.12 | 0.04 | 7.268 | 0.022 |  | 292.5429 | [5 empty] | 240 | Pb212 | Ra224 | Pb214 | 43.6 | 0.5 | 4.12 | 0.04 | 7.268 | 0.022 |  | 44.8023 | [5 empty] | 239.2078 | Pb212, Ra224, Pb214 | 292.5429 |  | 44.8023 |  | K40 | K-40 | 178.4296 | 17.5862
296.1614 | 0.29616 |  | 0.1221 | 0.1343 | 0.1079 | 0.042817 |  | 0.1223 | 0.13222 | 297 | Pb214 |  | 0.014991 | 0.0003078 | 0.015942 | 0.00025943 | [6 empty] | 0.022429 | 0.021806 | 0.023022 | 0.00060816 | 0.66837 | 0.022732 | 0.71075 | 0.022476 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.66837 | 0.91315 | 0.71075 | 0.91315 | 0 | 0.91287 | 0 | 0.91287 | 5.6651 | 6.4901 | 0.01 |  | 4.0265 | 0 | 5.1731 |  | 7.0664 | 8.3603 | 0.01 | 0.09856 | 0.09856 | 1.0469 | 0 | 0 | 0 | 8.4194 | 9.7004 | 0.01 | 0.82982 | 0 | 8.8146 |  | 297 | Pb214 |  |  | 18.414 | 0.036 |  |  |  |  | 28.0933 | 21.8664 |  | K-40 | 178.4296 |  |  | 297 | Pb214 |  |  | 18.414 | 0.036 |  |  |  |  | 47.8688 | 4.5064 |  | K-40 | 17.5862 |  |  | 296.1614 | Pb214 | 21.8664 | 28.0933 | 4.5064 | 47.8688 | Th232 | Th-232 | 1175.7312 | 119.9513
339.1264 | 0.33913 |  | 0.116 | 0.12759 | 0.1022 | 0.040556 |  | 0.1162 | 0.12562 | 340 | Ac228 |  | 0.0042818 | 0.0001645 | 0.003179 | 0.00011585 | [6 empty] | 0.019621 | 0.019066 | 0.020154 | 0.00054404 | 0.21823 | 0.01034 | 0.16203 | 0.0074194 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.21823 | 0.91293 | 0.16203 | 0.9129 | 0 | 0.91287 | 0 | 0.91287 | 5.6605 | 6.4839 | 0.01 |  | 0.91714 | 0 | 5.1675 |  | 7.0072 | 8.2785 | 0.01 | 0.087342 | 0.087342 | 0.28616 | 0 | 0 | 0 | 8.209 | 9.4184 | 0.01 | 0.71699 | 0 | 2.3491 |  | 340 | Ac228 |  |  | 11.4 | 0.4 |  |  |  |  | 45.3297 | 8.0451 | hmm… |  |  |  |  | 340 | Ac228 |  |  | 11.4 | 0.4 |  |  |  |  | 20.6071 | 6.2894 | hmm… |  |  |  |  | 339.1264 | Ac228 | 8.0451 | 45.3297 | 6.2894 | 20.6071 |  | Ac-228
353.115 | 0.35311 |  | 0.1142 | 0.12561 | 0.1006 | 0.039921 |  | 0.1144 | 0.12368 | 354 | Pb214 |  | 0.0203 | 0.00035818 | 0.023433 | 0.00031453 | [6 empty] | 0.018851 | 0.018316 | 0.019368 | 0.00052563 | 1.0768 | 0.035532 | 1.2431 | 0.038467 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 1.0768 | 0.91356 | 1.2431 | 0.91368 | 0 | 0.91287 | 0 | 0.91287 | 5.6592 | 6.4821 | 0.01 |  | 7.0347 | 0 | 5.1707 |  | 6.9895 | 8.2541 | 0.01 | 0.07037 | 0.07037 | 1.7848 | 0 | 0 | 0 | 8.1475 | 9.3361 | 0.01 | 0.57334 | 0 | 14.5416 |  | 354 | Pb214 |  |  | 35.6 | 0.7 |  |  |  |  | 14.5297 | 19.7604 |  | Th-232 | 1175.7312 |  |  | 354 | Pb214 |  |  | 35.6 | 0.7 |  |  |  |  | 40.8472 | 1.6105 |  | Th-232 | 119.9513 |  |  | 353.115 | Pb214 | 19.7604 | 14.5297 | 1.6105 | 40.8472 |  | Ra-224 | 266.6564 | 41.1548
511.9855 | 0.51199 |  | 0.09828 | 0.1081 | 0.0862 | 0.034206 |  | 0.09848 | 0.10646 | 513 | e-p |  | 0.016078 | 0.00031877 | 0.016142 | 0.00026105 | 0.013942 | 0.00041058 | 0.014568 | 0.00043884 | 0.013902 | 0.00023785 | 0.013037 | 0.012657 | 0.01341 | 0.0003764 | 1.2333 | 0.043196 | 1.2382 | 0.040977 | 1.0694 | 0.044106 | 1.1175 | 0.046628 | 1.0663 | 0.035788 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.97841 | 0.96085 | 1.0604 | 0.91674 | 0.97766 | 0.96055 | 1.0589 | 0.96055 | 0.25489 | 0.96182 | 0.17785 | 0.91765 | 0.09176 | 0.96157 | 0.058535 | 0.96169 | 5.6476 | 6.4662 | 0.01 |  | 1.0044 | 0.3785 | 5.1825 | 6.2185 | 6.8284 | 8.032 | 0.01 | 0.1078 | 0.1078 | 0.329 | 0.044636 | 0.044636 | 0.11341 | 7.6106 | 8.6257 | 0.01 | 0.82039 | 0.38501 | 2.5039 | 0.97821 | 513 | e-p | Tl208 |  | not very much |  | 22.5 | 0.2 |  |  |  | -0.0073987 |  | Ac-228 |  |  |  | 513 | e-p | Tl208 |  | much more |  | 22.5 | 0.2 |  |  |  | 0.77543 |  | Ac-228 |  |  |  | 511.9855 | e-p, Tl208 | No outright result as e-p has no branching ratio |  |  |  |  | Pb-212 | 5.073 | 0.58893
584.9261 | 0.58493 |  | 0.09281 | 0.10208 | 0.08135 | 0.032282 |  | 0.09301 | 0.10055 | 586 | Tl208 |  | 0.0065254 | 0.00020308 | 0.0075275 | 0.00017827 | [6 empty] | 0.011418 | 0.011083 | 0.011747 | 0.00033215 | 0.57152 | 0.024347 | 0.65929 | 0.024731 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.57152 | 0.9132 | 0.65929 | 0.91321 | 0 | 0.91287 | 0 | 0.91287 | 5.6437 | 6.4608 | 0.01 |  | 3.7208 | 0 | 5.1538 |  | 6.7709 | 7.953 | 0.01 | 0.021991 | 0.021991 | 0.96467 | 0 | 0 | 0 | 7.4296 | 8.3889 | 0.01 | 0.16339 | 0 | 7.1671 |  | 586 | Tl208 |  |  | 85 | 0.3 |  |  |  |  | 6.0633 | 4.3774 |  | Ra-224 | 266.6564 |  |  | 586 | Tl208 |  |  | 85 | 0.3 |  |  |  |  | 8.4318 | 0.19222 |  | Ra-224 | 41.1548 |  |  | 584.9261 | Tl208 | 4.3774 | 6.0633 | 0.19222 | 8.4318 |  | Bi-212 | 15.9239 | 3.2147
610.9049 | 0.6109 |  | 0.09106 | 0.10016 | 0.0798 | 0.031667 |  | 0.09126 | 0.098659 | 612 | Bi214 |  | 0.015354 | 0.00031151 | 0.01801 | 0.00027575 | [6 empty] | 0.010934 | 0.010613 | 0.011251 | 0.00031873 | 1.4043 | 0.049874 | 1.6472 | 0.054237 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 1.4043 | 0.91423 | 1.6472 | 0.91448 | 0 | 0.91287 | 0 | 0.91287 | 5.6424 | 6.459 | 0.01 |  | 9.2942 | 0 | 5.1599 |  | 6.7523 | 7.9275 | 0.01 | 0.027848 | 0.027848 | 2.4003 | 0 | 0 | 0 | 7.372 | 8.314 | 0.01 | 0.2053 | 0 | 17.6947 |  | 612 | Bi214 |  |  | 45.49 | 0.19 |  |  |  |  | 11.3433 | 20.4312 |  | Pb-212 | 5.073 |  |  | 612 | Bi214 |  |  | 45.49 | 0.19 |  |  |  |  | 38.898 | 0.4513 |  | Pb-212 | 0.58893 |  |  | 610.9049 | Bi214 | 20.4312 | 11.3433 | 0.4513 | 38.898 |  | Tl-208 | 4.4969 | 0.19985
728.8088 | 0.72881 |  | 0.08415 | 0.092555 | 0.07368 | 0.029238 |  | 0.08434 | 0.091178 | 730 | Bi212 |  | 0.0017285 | 0.00010452 | 0.0017225 | 8.5277e-05 | [6 empty] | 0.0091701 | 0.0088996 | 0.0094379 | 0.00026914 | 0.18849 | 0.012669 | 0.18784 | 0.010811 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.18849 | 0.91296 | 0.18784 | 0.91293 | 0 | 0.91287 | 0 | 0.91287 | 5.6375 | 6.4521 | 0.01 |  | 1.0589 | 0 | 5.1466 |  | 6.6776 | 7.8251 | 0.01 | 0.029914 | 0.029914 | 0.29296 | 0 | 0 | 0 | 7.1465 | 8.0217 | 0.01 | 0.21378 | 0 | 2.0937 |  | 730 | Bi212 |  |  | 6.65 | 0.04 |  |  |  |  | 77.3933 | 15.9239 |  | Bi-212 | 15.9239 | B.R. | expected | 730 | Bi212 |  |  | 6.65 | 0.04 |  |  |  |  | 31.4835 | 3.2147 |  | Bi-212 | 3.2147 | B.R. | expected | 728.8088 | Bi212 | 15.9239 | 77.3933 | 3.2147 | 31.4835 |  | e-p | -0.0073987 | 0.77543
768.7762 | 0.76878 |  | 0.08211 | 0.090311 | 0.07189 | 0.028528 |  | 0.08229 | 0.088962 | 770 | Bi214 |  | 0.0019181 | 0.0001101 | 0.0024529 | 0.00010176 | [6 empty] | 0.0086945 | 0.0084378 | 0.008949 | 0.0002556 | 0.22061 | 0.014228 | 0.28212 | 0.014345 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.22061 | 0.91298 | 0.28212 | 0.91298 | 0 | 0.91287 | 0 | 0.91287 | 5.636 | 6.4501 | 0.01 |  | 1.59 | 0 | 5.1456 |  | 6.6552 | 7.7945 | 0.01 | -0.018304 | 0 | 0.40352 | 0 | 0 | 0 | 7.0805 | 7.9366 | 0.01 | 0 | 0 |  |  | 770 | Bi214 |  |  | 4.892 | 0.016 |  |  |  |  | 105.1841 | 32.5026 |  | Tl-208 | 4.4969 | 35.93(7)% | 5.7215 | 770 | Bi214 |  |  | 4.892 | 0.016 | [5 empty] | 0 |  | Tl-208 | 0.19985 | 35.93(7)% | 1.155 | 768.7762 | Bi214 | 32.5026 | 105.1841
913.6582 | 0.91366 |  | 0.07571 | 0.083272 | 0.06625 | 0.02629 |  | 0.07583 | 0.081978 | 915 | Ac228 |  | 0.004484 | 0.00016834 | 0.0046609 | 0.00014028 | [6 empty] | 0.0073184 | 0.0071017 | 0.0075338 | 0.00021607 | 0.6127 | 0.029264 | 0.63687 | 0.026851 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.6127 | 0.91334 | 0.63687 | 0.91327 | 0 | 0.91287 | 0 | 0.91287 | 5.6315 | 6.4436 | 0.01 |  | 3.5865 | 0 | 5.143 |  | 6.5837 | 7.6969 | 0.01 | 0.067943 | 0.067943 | 0.98434 | 0 | 0 | 0 | 6.8749 | 7.6727 | 0.01 | 0.4671 | 0 | 6.7673 |  | 915 | Ac228 |  |  | 26.2 | 0.8 |  |  |  |  | 19.6344 | 13.689 | hmm… |  |  |  |  | 915 | Ac228 |  |  | 26.2 | 0.8 |  |  |  |  | 25.8294 | 1.7828 | hmm… |  |  |  |  | 913.6582 | Ac228 | 13.689 | 19.6344 | 1.7828 | 25.8294
971.6109 | 0.97161 |  | 0.0735 | 0.080841 | 0.06431 | 0.02552 |  | 0.07363 | 0.0796 | 973 | Ac228 |  | 0.0034634 | 0.00014795 | 0.0039094 | 0.00012847 | [6 empty] | 0.0068826 | 0.0066786 | 0.0070855 | 0.00020346 | 0.5032 | 0.026141 | 0.56801 | 0.025107 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.5032 | 0.91325 | 0.56801 | 0.91322 | 0 | 0.91287 | 0 | 0.91287 | 5.6299 | 6.4414 | 0.01 |  | 3.1979 | 0 | 5.1413 |  | 6.5586 | 7.6628 | 0.01 | 0.015618 | 0.015618 | 0.85363 | 0 | 0 | 0 | 6.8045 | 7.5829 | 0.01 | 0.10627 | 0 | 5.8085 |  | 973 | Ac228 |  |  | 15.9 | 0.5 |  |  |  |  | 32.3416 | 20.1123 | hmm… | e-p | -0.0073987 | next to nothing |  | 973 | Ac228 |  |  | 15.9 | 0.5 |  |  |  |  | 36.5316 | 0.66839 | hmm… | e-p | 0.77543 | oddly high |  | 971.6109 | Ac228 | 20.1123 | 32.3416 | 0.66839 | 36.5316
1122.488 | 1.1225 |  | 0.06845 | 0.075287 | 0.05987 | 0.023758 |  | 0.06863 | 0.074195 | 1124 | Bi214 |  | 0.002942 | 0.00013636 | 0.0036139 | 0.00012352 | [6 empty] | 0.0059588 | 0.0057818 | 0.0061349 | 0.00017656 | 0.49372 | 0.02716 | 0.60648 | 0.027434 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.49372 | 0.91327 | 0.60648 | 0.91328 | 0 | 0.91287 | 0 | 0.91287 | 5.6263 | 6.4365 | 0.01 |  | 3.4123 | 0 | 5.1384 |  | 6.5004 | 7.5838 | 0.01 | -0.031213 | 0 | 0.89155 | 0 | 0 | 0 | 6.6447 | 7.3798 | 0.01 | 0 | 0 |  |  | 1124 | Bi214 |  |  | 14.91 | 0.03 |  |  |  |  | 34.4631 | 22.8859 | [5 empty] | 1124 | Bi214 |  |  | 14.91 | 0.03 | [5 empty] | 0 | [5 empty] | 1122.488 | Bi214 | 22.8859 | 34.4631
1242.3903 | 1.2424 |  | 0.06501 | 0.071503 | 0.05687 | 0.022567 |  | 0.06515 | 0.070432 | 1244 | Bi214 |  |  |  | 0.0021215 | 9.4639e-05 | [6 empty] | 0.0053843 | 0.0052243 | 0.0055438 | 0.00015975 | 0 |  | 0.39401 | 0.021109 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.39401 | 0.91311 | 0 | 0.91287 | 0 | 0.91287 | 5.6239 | 6.433 | 0.01 |  | 2.2159 | 0 | 5.1353 |  | 6.4601 | 7.5292 | 0.01 | -0.34301 | 0 | 0.48509 | 0 | 0 | 0 | 6.5367 | 7.2432 | 0.01 | 0 | 0 |  |  | 1244 | Bi214 |  |  | 5.831 | 0.014 |  |  |  |  | 88.0689 | 38.0017 | hmm… |  |  |  |  | 1244 | Bi214 |  |  | 5.831 | 0.014 | [5 empty] | 0 | [5 empty] | 1242.3903 | Bi214 | 38.0017 | 88.0689
1464.2095 | 1.4642 |  | 0.05972 | 0.065685 | 0.05228 | 0.020746 |  | 0.05985 | 0.064703 | 1466 | K40 |  | 0.018928 | 0.00034587 | 0.019691 | 0.00028832 | 0.0045041 | 0.00023337 | 0.0046004 | 0.00024661 | 0.0044012 | 0.00013383 | 0.0045693 | 0.0044334 | 0.0047049 | 0.00013579 | 4.1425 | 0.14452 | 4.3093 | 0.14277 | 0.98573 | 0.058878 | 1.0068 | 0.061709 | 0.9632 | 0.040954 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.90869 | 0.9693 | 0.95991 | 0.91646 | 0.90815 | 0.96831 | 0.9591 | 0.96831 | 3.2338 | 0.98001 | 3.3494 | 0.92751 | 0.077578 | 0.9701 | 0.047703 | 0.97027 | 5.6202 | 6.4277 | 0.01 |  | 18.8243 | 0.30662 | 5.2129 | 6.2366 | 6.3971 | 7.444 | 0.01 | 0.29117 | 0.29117 | 5.2702 | 0.036388 | 0.036388 | 0.097013 | 6.3721 | 7.0361 | 0.01 | 1.8553 | 0.25603 | 33.5823 | 0.68259 | 1466 | K40 |  |  | 10.55 | 0.11 |  |  |  |  | 49.4468 | 178.4296 |  |  | 1852.6978 |  |  | 1466 | K40 |  |  | 10.55 | 0.11 |  |  |  |  | 318.3157 | 17.5862 |  |  | 216.6189 |  |  | 1464.2095 | K40 | 178.4296 | 49.4468 | 17.5862 | 318.3157
1767.962 | 1.768 |  | 0.0541 | 0.059504 | 0.04746 | 0.018833 |  | 0.05414 | 0.05853 | 1770 | Bi214 |  | 0.0027903 | 0.0001328 | 0.003312 | 0.00011825 |  |  |  |  | 0.00063892 | 5.0991e-05 | 0.0037848 | 0.0036721 | 0.0038973 | 0.00011263 | 0.73723 | 0.041381 | 0.87509 | 0.040673 | 0 |  | 0 |  | 0.16881 | 0.014379 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.16 | 0.92341 | 0.16829 | 0.91349 | 0.15991 | 0.9232 | 0.16816 | 0.9232 | 0.57724 | 0.92433 | 0.7068 | 0.9144 | -0.15991 | 0.9232 | -0.16816 | 0.9232 | 5.6164 | 6.422 | 0.01 |  | 3.9696 | 0 | 5.1356 |  | 6.3288 | 7.3518 | 0.01 | -0.050003 | 0 | 1.0583 | -0.15991 | 0 | 0.15991 | 6.1991 | 6.8198 | 0.01 | 0 | 0 |  |  | 1770 | Bi214 |  |  | 15.31 | 0.05 |  |  |  |  | 33.5442 | 25.9284 | [5 empty] | 1770 | Bi214 |  |  | 15.31 | 0.05 | [5 empty] | 0 | [5 empty] | 1767.962 | Bi214 | 25.9284 | 33.5442
2207.6037 | 2.2076 |  | 0.04793 | 0.052717 | 0.04231 | 0.01679 |  | 0.04798 | 0.05187 | 2210 | Bi214 |  |  |  | 0.00076626 | 5.6877e-05 | [6 empty] | 0.0030314 | 0.002941 | 0.0031217 | 9.0308e-05 | 0 |  | 0.25277 | 0.020217 | 0 |  | 0 |  | 0 |  | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0 | 0.91287 | 0.25277 | 0.91309 | 0 | 0.91287 | 0 | 0.91287 | 5.6122 | 6.4159 | 0.01 |  | 1.4186 | 0 | 5.1245 |  | 6.252 | 7.2486 | 0.01 | -0.22691 | 0 | 0.3209 | 0 | 0 | 0 | 6.0116 | 6.5869 | 0.01 | 0 | 0 |  |  | 2210 | Bi214 |  |  | 4.913 | 0.023 |  |  |  |  | 104.3046 | 28.8748 | [5 empty] | 2210 | Bi214 |  |  | 4.913 | 0.023 | [5 empty] | 0 | [5 empty] | 2207.6037 | Bi214 | 28.8748 | 104.3046
2621.2666 | 2.6213 |  | 0.04366 | 0.048021 | 0.03878 | 0.015389 |  | 0.04362 | 0.047157 | 2624 | Tl208 |  | 0.0039563 | 0.00015813 | 0.0041501 | 0.00013237 | 0.0019226 | 0.00015247 | 0.0020821 | 0.00016591 | 0.0020592 | 9.1542e-05 | 0.0025532 | 0.0024771 | 0.0026293 | 7.6105e-05 | 1.5496 | 0.077259 | 1.6254 | 0.070959 | 0.753 | 0.063796 | 0.81548 | 0.069377 | 0.80651 | 0.043167 | 1 | 0.2 | 6 | 0.4 | 0.4 | 0.1 | 0.77132 | 0.96581 | 0.80446 | 0.91593 | 0.77098 | 0.96468 | 0.804 | 0.96468 | 0.77823 | 0.9689 | 0.82097 | 0.91868 | -0.017978 | 0.96679 | 0.011483 | 0.96717 | 5.6094 | 6.4116 | 0.01 |  | 4.6052 | 0.073627 | 5.1532 | 6.2011 | 6.1977 | 7.1758 | 0.01 | 0.035179 | 0.035179 | 1.3076 | -0.028238 | 0 | 0.023103 | 5.8833 | 6.4283 | 0.01 | 0.20697 | 0 | 7.6931 |  | 2624 | Tl208 |  |  | 99.755 | 0.004 |  |  |  |  | 5.1659 | 4.6165 | [5 empty] | 2624 | Tl208 |  |  | 99.755 | 0.004 |  |  |  |  | 7.712 | 0.20748 | [5 empty] | 2621.2666 | Tl208 | 4.6165 | 5.1659 | 0.20748 | 7.712
|  | 0.90919 | CH5OH | H2O | density | 2.52 | g/cm^3 | [54 empty] | Bq | 83.3959 | 0.80498 | 1.1051 | 3.0847 |  |  | Bq |  | 1.1107 |  |  | 0.15226 |  |  |  | Bq | 8.6975 | 1.4617 | 2.8112 | 0.26603
|  | % | 0.431 | 0.569 | SiO2 | 74 | [55 empty] | kg | 0.64057 | 0.09432 | 0.001 | 1e-05 |  |  | kg |  | 0.63643 |  |  | 0.63643 |  |  |  | kg | 0.63643 | 0.63643 | 0.01 | 0.01
|  | Density | 0.7893 | 1 | Na2O | 13
[5 empty] | CaO | 10.5 | [55 empty] | Bq/kg | 130.1905 | 8.5345 | 1.7372 | 32.7043 |  |  | Bq/kg | Bq/kg | 1.7453 |  |  | 0.23925 |  |  |  | Bq/kg | 13.666 | 2.2967 | 4.4223 | 0.41955
[5 empty] | Al2O3 | 1.3
[5 empty] | K2O | 0.3
[5 empty] | MgO | 0.2
[5 empty] | Fe2O3 | 0.04
[5 empty] | TiO2 | 0.01
[5 empty] | SO3 | 0.2
https://en.wikipedia.org/wiki/Soda%E2%80%93lime_glass

## Data
Isotope Energy (keV) Branching ratio (%) | seen? (glenkinchie) | [10 empty] | foil |  |  |  |  | glenkinchie |  |  | 70cl |  | highland park
|  |  | kBq | unc | Year | unc | Half Life/y | unc | Current Activity/kBq | unc |  | approx by eye |  |  |  |  | approx by eye |  |  | both |  | approx by eye
212Pb 238.63 43.60 | x
212Pb 300.09 3.30 |  | Cs137 | 39.89 | 2.3934 | 1977.3333 | 0.5 | 30.08 | 0.09 | 15.2714 | 0.93406 | maybe??? | 2665 | Sc44 + Sc44? |  | 2662.2332 |  | 46 | Pb210 | 45.3658 |  |  | 47 | Pb210 | 46.365
214Pb 242.00 7.25 |  | Eu152 | 37 | 0.5 | 1981 | 0.5 | 13.517 | 0.009 | 5.2714 | 0.15293 |  | 2625 | Tl208 |  | 2622.2658 |  | 63 | Th234 | 62.3519 |  |  | 64 | Th234 | 63.3511
214Pb 295.22 18.42 | x | Co60 | 42.8 | 1.284 | 2008 | 0.5 | 5.2711 | 0.0003833 | 10.0749 | 0.72812 | maybe??? | 2111 | Tl208 - e-p? |  | 2108.6843 |  | 77 |  | 76.3405 |  |  | 77 |  | 76.3405
214Pb 351.93 35.60 | x | [9 empty] | maybe??? | 1770 | Bi214 | ? | 1767.962 |  | 93 | Th234 | 92.3275 |  |  | 93 | Th234 | 92.3275
212Bi 727.33 6.67 | [5 empty] | days | 1925.28 | 0.14 |  |  |  | 1673.5 | Sc44 + e-p |  | 1671.5405 | [5 empty] | ??? | 145 |  | 144.2852
212Bi 785.37 1.10 | [11 empty] | 1505 | Sc44 |  | 1503.1777 |  | 187 | Ra226 | 186.251 |  |  | 187 | Ra226 | 186.251
212Bi 1620.50 1.47 |  | Na22 | 3340 | 133.6 | 1996 | 0.5 | 2.6018 | 0.0022 | 7.2884 | 1.0144 |  | 1465 | K40 |  | 1463.2103 |  | 240 | Pb212 | 239.2078 |  |  | 240 | Pb212 | 239.2078
214Bi 609.32 45.49 | x | Ba133 | 45.5 | 2.275 | 1995 | 0.5 | 10.551 | 0.11 | 9.403 | 0.58338 |  | 1160 | Sc44 |  | 1158.4587 | [5 empty] | ??? | 275 |  | 274.1793
214Bi 1120.29 14.92 | x | Bi207 | 111 | 0.5 | 1982 | 0.0027379 | 31.55 | 0.04 | 49.2371 | 0.22754 |  | 511 | e-p |  | 509.9871 |  | 296 | Pb214 | 295.1622 |  |  | 297 | Pb214 | 296.1614
214Bi 1764.49 15.30 | x | [10 empty] | 145 | Ti44/additon of Ti44 |  | 144.2852 | ??? | 339 | Ac228 | 338.1272 |  |  | 341 | Ac228 | 340.1256
234Th 63.29 3.70 | x | [6 empty] | https://www.nndc.bnl.gov/nudat2/ |  |  | tiny | 92 | Th234 |  | 91.3283 |  | 353 | Pb214 | 352.1158 |  |  | 354 | Pb214 | 353.115
234Th 92.38 2.13 | x | [10 empty] | 79 | Ti44 |  | 78.3389 |  | 512 | e-p | 510.9863 |  |  | 513 | e-p | 511.9855
234Th 92.80 2.10 | x | Ba133-1 | 42.37 | 2.9659 | 1977.3333 | 0.0027379 | 10.551 | 0.011 | 2.7433 | 0.19219 |  | 68 | Ti44 |  | 67.3479 |  | 585 | Tl208 | 583.9269 |  |  | 586 | Tl208 | 584.9261
208Tl 583.19 85.00 | x | Co60-1 | 185 | 0.5 | 1960 | 0.5 | 5.2711 | 0.0003833 | 0.079018 | 0.0052 | tiny | 46 | Pb210 |  | 45.3658 |  | 611 | Bi214 | 609.9057 |  |  | 612 | Bi214 | 610.9049
208Tl 860.56 12.50 |  | Bi207-2 | 102.8 | 3.7522 | 1996.5 | 0.5 | 31.55 | 0.04 | 62.7065 | 2.3905 | [11 empty] | 730 | Bi212??? | 728.8088
208Tl 2614.51 99.75 | x | [20 empty] | 770 | Bi214 | 768.7762
228Ac 338.32 11.27 |  | Am241 | 4.625 | 0.4625 | 1960 | 0.5 | 432.6 | 0.6 | 4.2078 | 0.42079 | [5 empty] | tiny | 914 | Ac228 | 912.659 |  |  | 915 | Ac228 | 913.6582
228Ac 911.20 25.80 | x | Am241-2 | 38.1 | 1.524 | 1994.75 | 0.5 | 432.6 | 0.6 | 36.648 | 1.4662 | [11 empty] | 973 | Ac228 | 971.6109
228Ac 968.97 15.80 | [20 empty] | tiny | 1005 |  | 1003.5849
[16 empty] | tiny | 1123 | Bi214 | 1121.4888 |  | ??? | 1124 | Bi214 | 1122.488
[16 empty] | ??? | 1241 | Bi214 | 1239.3927 |  |  | 1242 | Bi214 | 1240.3919
[17 empty] | 1464 | K40 | 1462.2111 |  |  | 1466 | K40 | 1464.2095
[21 empty] | tiny | 1595 |  | 1593.1045
[16 empty] | tiny | 1768 | Bi214 | 1765.9636 |  |  | 1771 | Bi214 | 1768.9611
[21 empty] | ??? | 2110 |  | 2107.6851
[21 empty] | ??? | 2127 |  | 2124.6713
[21 empty] | tiny | 2213 |  | 2210.6012
[17 empty] | 2619 | Tl208 | 2616.2706 |  |  | 2624 | Tl208 | 2621.2666

## Glenkinchie(12yo)
peak | mean | std dev | amp |  | energy corrected | isotope | area | background area | difference | energy efficiency | en_eff_corrected | geometrical efficiency | Bq/cm^3 | Bq/kg
6 | 63.517 | 0.505 | 0.002211 |  | 63.0122 | Th234 | 0.0027988 | 0.0076455 | -0.0048468 | 0.090024 | 0.031089 | 1 | 0.031089
9 | 93.143 | 0.87 | 0.0020119 |  | 92.6134 | Th234 | 0.0043874 | 0.014868 | -0.01048 | 0.067214 | 0.065275 | 1 | 0.065275
22 | 186.529 | 0.487 | 0.0017682 |  | 185.9211 | Ra226 | 0.0021585 | 0.0047264 | -0.002568 | 0.035371 | 0.061023 | 1 | 0.061023
29 | 239.514 | 0.407 | 0.0018475 |  | 238.8617 | Pb212 | 0.0018848 |  | 0.0018848 | 0.027721 | 0.067991 | 1 | 0.067991
36 | 296.452 | 0.695 | 0.00096166 |  | 295.752 | Pb214 | 0.0016753 |  | 0.0016753 | 0.02247 | 0.074557 | 1 | 0.074557
44 | 353.15 | 0.986 | 0.0011419 |  | 352.4025 | Pb214 | 0.0028223 |  | 0.0028223 | 0.018896 | 0.14936 | 1 | 0.14936
60 | 512.309 | 1.543 | 0.00096075 |  | 511.4281 | e-p | 0.0037159 | 0.011581 | -0.0078651 | 0.013054 | 0.28466 | 1 | 0.28466
62 | 584.493 | 0.388 | 0.00046673 |  | 583.5516 | Tl208 | 0.00045393 |  | 0.00045393 | 0.011447 | 0.039655 | 1 | 0.039655
64 | 611.032 | 1.189 | 0.00072557 |  | 610.0684 | Bi214 | 0.0021625 |  | 0.0021625 | 0.010951 | 0.19746 | 1 | 0.19746
70 | 913.534 | 0.423 | 0.00027122 |  | 912.3168 | Ac228 | 0.00028758 |  | 0.00028758 | 0.0073302 | 0.039232 | 1 | 0.039232
73 | 1123.265 | 0.91 | 0.00025538 |  | 1121.8721 | Bi214 | 0.00058254 |  | 0.00058254 | 0.0059627 | 0.097698 | 1 | 0.097698
75 | 1464.02 | 1.524 | 0.00040207 |  | 1462.3415 | K40 | 0.001536 | 0.0037798 | -0.0022439 | 0.0045755 | 0.33569 | 1 | 0.33569
76 | 1768.074 | 1.511 | 0.00014443 |  | 1766.1406 | Bi214 | 0.00054705 | 0.00059626 | -4.9217e-05 | 0.0037889 | 0.14438 | 1 | 0.14438
77 | 2619.719 | 2.157 | 0.00012951 |  | 2617.0719 | Tl208 | 0.00070024 | 0.0015466 | -0.00084631 | 0.0025574 | 0.27381 | 1 | 0.27381
https://www.osti.gov/servlets/purl/4331753
The nuclear weapons archive
|  |  |  | https://www.nndc.bnl.gov/nudat2/reCenter.jsp?z=81&n=119 | [5 empty] | https://en.wikipedia.org/wiki/File:Decay_Chain_Thorium.svg
|  |  |  | http://www.nucleide.org/Laraweb/index.php | [5 empty] | https://en.wikipedia.org/wiki/File:Decay_chain(4n%2B2,_Uranium_series).svg
http://www.nucleide.org/DDEP_WG/DDEPdata.htm

## Overall Background
peak | mean | std dev | amp | energy corrected | isotope | area
4 | 46.532 | 0.467 | 0.0045516 | 46.0415 | Pb-210 | 0.0053281
6 | 63.549 | 0.79 | 0.0038609 | 63.0442 | Th234 | 0.0076455
9 | 93.045 | 0.93 | 0.0063778 | 92.5155 | Th234 | 0.014868
18 | 186.505 | 0.442 | 0.004266 | 185.8972 | Ra226 | 0.0047264
38 | 512.276 | 1.614 | 0.0028626 | 511.3951 | e-p | 0.011581
51 | 1463.406 | 1.854 | 0.00081334 | 1461.728 | K40 | 0.0037798
52 | 1767.717 | 1.093 | 0.00021763 | 1765.7839 | Bi214 | 0.00059626
53 | 2618.939 | 2.511 | 0.00024571 | 2616.2925 | Tl208 | 0.0015466

## Decay Chains
isotopes with gamma emissions of note |  | (>40keV) | (>P/100=1)
Uranium Series |  | Thorium Series
Th-234 |  | Ac-228
Ra-226 |  | Th-228
Pb-214 |  | Ra-224
Bi-214 |  | Pb-212
Tl-210 |  | Bi-212 | (just)
Pb-210 | single very low energy | Tl-208
Hg-206